## 표지(영문)
SCOPE OF ACCREDITATION TO ISO/IEC 17025:2017 & KS Q ISO/IEC 17025:2017
SS Technology Inc.
61, Magokjungang 8-ro 3-gil, Gangseo-gu, Seoul, Republic of Korea
Phone : 02-2613-3061 , Fax : 02-3667-3063 , e-mail: cal@sstech.co.kr
CALIBRATION
Valid To : February. 22, 2027. | [5 empty] | Accreditation No : KC07-210
In recognition of the successful completion of the KOLAS evaluation process,
accreditation is granted to this laboratory to perform the following calibrations
Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site
102. Linear dimension |  |  | 201. Mass |  |  | 402. Resistance,
10201 | Balls | N | 20109 | Electric balances | Y | Capacitance and Inductance
10209 | End bars | N | 20113 | Spring scale balances | Y | 40201 | Capacitance bridges /indicators | Y
10210 | Extensometers, | Y | 20116 | Weights | N | 40202 | Decade capacitors | Y
| linear displacement transducers |  | 202. Force |  |  | 40205 | Earth testers | Y
10211 | Filler gauges | Y | 20203 | Tension/compression | Y | 40206 | Inductance bridges /indicators | Y
10213 | Gap gauges | N |  | testing machines |  | 40208 | Inductors | Y
10216 | Height gauges/measuring machines | Y | 20204 | Push-pull gauges | Y | 40210 | Insulation testers | Y
10219 | Linear scale | N | 203. Torque |  |  | 40213 | Resistance bridges & | N
10220 | Standard measuring machines | Y | 20303 | Torque wrenches/drivers | Y |  | similar instruments
10223 | Electronic micrometers | Y | 204. Pressure |  |  | 40214 | Resistance meters | Y
10225 | Laser scan micrometer | Y | 20402 | Manometers | N | 40215 | Resistors | Y
10228 | Cylindrical plug/pin gauges, | Y | 20403 | Penumatic Pressure Balance | N | 40217 | Impedance bridges/LCR meters | Y
| thread measuring wire gauges |  | 20404 | Hydraulic Pressure Balance | N | 403. AC voltage, current & power
10229 | Radius gauges | N | 20406 | Absolute pressure gauges | Y | 40301 | AC ammeters | Y
10230 | Cylindrical ring gauges | N | 20408 | Compound pressure gauges | Y | 40302 | Clamp ammeters/voltmeters | Y
10232 | Step gauges/Caliper testers | N | 20409 | Differential pressure gauges | Y | 40303 | AC voltage/current calibrators | Y
10233 | Taper thickness gauges | N | 20411 | Gauge pressure gauges | Y | 40304 | Wattmeter calibrators | N
10234 | Ultrasonic thickness gauges | Y | 20412 | Pressure transducers/ | Y | 40305 | AC current shunts
10235 | Ultrasonic | N |  | transmitters |  | 40307 | Voltage current phase meter | Y
| /coating thickness specimens |  | 20413 | Dial type vacuum gauges | Y | 40310 | Power factor meters | Y
| /Ultrasonic thickness specimens |  | 206. Volume |  |  | 40311 | AC power meters | Y
10236 | Coating thickness testers | Y | 20601 | Volumetric glasswares | N | 40312 | AC power supplies | Y
103. Angle |  |  | 20606 | Piston type volume meters | N | 40313 | Puncture/safety testers | Y
10304 | Bevel protractors | N | 301. Time/frequency |  |  | 40318 | AC voltmeters | Y
10311 | Plate/Square/Electric levels | N | 30102 | Frequency standards | N | 404. Other DC & LF Measurements
104. From |  |  | 30103 | General frequency sources | Y | 40401 | LF amplifiers | Y
10407 | Precision surface plates | Y | 30104 | Frequency meters/counters | Y | 40402 | DC/LF attenuators | Y
105. Complex geometry |  |  | 30106 | Time interval meters/ | Y | 40403 | Multimeter calibrators | N
10503 | Contact coordinate | Y |  | Stop watches/Timers | Y | 40404 | Oscilloscope calibrators | N
| measuring machines |  | 302. Velocity & revolution |  |  | 40407 | Audio distortion analyzers | Y
10504 | Non-contact coordinate | Y | 30201 | Standard RPM generators | Y |  | /meters
| measuring machines |  | 30202 | Contact type tachometers | Y | 40408 | LF filters | Y
10511 | Measuring microscopes, | Y | 30203 | Photo tachometers/stroboscopes | Y | 40409 | LF/Audiosignalanalyzers | Y
| profile projectors |  | 401. DC Voltage & current |  |  | 40410 | Line frequency meters | Y
10512 | Micro measuring microscopes | Y | 40101 | DC ammeters | Y | 40411 | Function generators | Y
10517 | Stylus Type Roughness Testers | Y | 40102 | Transconductance amplifiers | N | 40413 | AC/DC high voltages meters | Y
10519 | Roughness Standard/comparison specimens | N | 40103 | DC voltage/current calibrators | Y | 40416 | Leakage current testers | Y
10525 | Thread plug gauges | N | 40104 | Electrical temperature calibrators | Y | 40417 | Electronic AC/DC loads | Y
106. Various dimensional |  |  | 40105 | DC current shunts | Y | 40419 | Analogue/Digital multimeters | Y
10601 | Inside/outside/gear tooth | Y | 40108 | DC power supplies | Y | 40420 | Noise meters | Y
| calipers, caliper gauges |  | 40112 | DC voltmeters | Y | 40421 | Oscilloscopes | Y
10603 | Cylinder/bore gauges | Y | 40113 | Static/Ionicvoltmeters | N | 40424 | Volt/Current recorders | Y
10604 | Depth gauges, | Y |  |  |  | 40425 | Relay test sets | Y
| depth micrometers |  |  |  |  | 40426 | LF signal generators | Y
10605 | Dial/digital gauges | Y |  |  |  | 40432 | Transistor curve tracers | Y
10609 | Micro indicators, | Y |  |  |  | 40434 | AC/DC high voltage generators | Y
| test indicators |  |  |  |  | 40435 | AC/DC high voltage probes | N
10610 | Micrometer heads | Y |  |  |  | 40436 | Logic analyzers | Y
10611 | 3-point micrometers | Y
10612 | Inside micrometers | Y
10613 | Outside micrometers | Y
10617 | Test Seive | N
Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site
501. Contact thermometry |  |  | 704. Fiber optics
50101 | Temperature generators: | Y | 70402 | Broadband light sources | Y
| ovens, furnaces, isothermal |  | 70410 | Optical attenuators | Y
| liquid baths, ice-point baths, |  | 70412 | Fiber-optic power meters | Y
| dry-block calibrators |  | 70413 | Optical loss testers | Y
50102 | Temperature indicators | Y | 70415 | Optical multimeters | Y
| /Recorders/Controllers, |  | 70417 | Optical spectrum analyzers | Y
| Temperature calibrators |  | 70418 | Optical time domain | Y
| Temperature calibrators |  |  | reflectometers;OTDR
50103 | Glass thermometers; | N | 70426 | Multi-laser wavelength meters | Y
| liquid-in-glass, Beckmann |  | 70429 | Frequency stabilized lasers | Y
50104 | Resistance thermometers; | Y |  | and LDs
| SPRT, IPRT, thermistors,etc. |  | 70430 | ASE light sources | Y
50105 | Thermal expansion | Y | 70431 | CW-laser wavelength meters | Y
| thermometers ; bimetal, |  | 70433 | Optical power stabilized | Y
| gas or liquid type |  |  | and LDs
50106 | Thermomecouples: noble metal, | Y | 901. Chemical analysis
| base metal, pure metal, |  | 90103 | Gas analyzers | Y
special type, etc.
50107 | Temperature transducers | Y
502. Non Contact thermometry
50204 | Radiation thermometers | N
50205 | Thermal image apparatus | N
50206 | Blackbody furnaces | Y
50207 | Others Thermometer; | N
ear thermometers, etc.
503. Humidity
50302 | Relative humidity hygrometers; | Y
polimer thinfilm, hair, etc.
50303 | Psychrometers; assmann | N
ventilated, PRT type, etc.
50304 | Temperature humidity recorders; | N
Hygrothermograph, etc
50305 | Dew-point/Relative humidity | N
Transducers
50306 | Humidity generators; | Y
two-pressure,
two-temperature, flow mixing
humidity gererator,
constant temperature and
humidity chamber, etc.
601. Sound in air
60106 | Sound level meters | N
603. Vibration
60301 | Vibration calibrators | N
60302 | Vibration transducers | N
60303 | Vibration measuring instruments | N
701. Photometry
70101 | Illuminance meters | N
Note
1. This laboratory provides calibration services in permanent standard laboratory and at on-site.
2. Laboratory conducts on-site calibration should meet requirements of KOLAS-SR-007.
3. On-site calibration is allowed to items with marking 'Y', not allowed to items with marking 'N'.
4. Measurement uncertainty normally is quoted as an expanded uncertainty at a coverage probability of 95 %, which usually requires the use of a coverage factor of K=2. It expresses the lowest uncertainty of measurement that can be provided by accredited calbration laboratories in normal conditions.
5.Due to the calibration environment such as reference standards or customers’ facilities, it is note that uncertainty of measurement on a calibration certificate may be expressed larger than measurement uncertainty on scope of accreditation in general.
6. If continuous calibration range is divided, each divided range`s endpoint indicates inclusive. * ex) If calibration range is divided to (0 ~ 25) mm and (25 ~ 100) mm, 25 mm in first range indicates inclusive and 25 mm in second range indicates exclusive.

## 교정측정능력(영문)
102. Linear dimension
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Balls | 10201 |  |  | Gauge blocks,
Standard measuring machines/
|  | (0.4 ~ 20) mm | 0.6 μm | SS-CAL-10201
End bars | 10209 |  |  | Gauge blocks,
Linear height gauge/
|  | (25 ~ 600) mm |  | SS-CAL-10209
Extensometers, | 10210 |  |  | Gauge blocks/
linear displacement transducers |  | (0 ~ 100) mm |  | SS-CAL-10210
(100 ~ 1 000) mm
Filler gauges | 10211 |  |  | Outside Micrometer/
|  | (0 ~ 5) mm | 1.4 μm | SS-CAL-10211
Gap gauges | 10213 |  |  | Gauge blocks, Contact coordinate
measuring machines/
|  | (4 ~ 100) mm | 1.5 μm | SS-CAL-10213
Height gauges/measuring machines | 10216 |  |  | Step gauge/
|  | (0 ~ 1 000) mm |  | SS-CAL-10216
Linear scale | 10219 |  |  | Gauge blocks/
|  | (0 ~ 1 000) mm |  | SS-CAL-10219
Standard measuring machines | 10220 |  |  | Gauge blocks /
|  | (0 ~ 500) mm |  | SS-CAL-10220
Electronic micrometers | 10223 |  |  | Gauge blocks/
|  | (0 ~ 5) mm | 0.22 μm | SS-CAL-10223
Laser scan micrometer | 10225 |  |  | Pin gauge/
|  | (0 ~ 25) mm | 0.6 μm | SS-CAL-10225
Cylindrical plug/pin gauges, | 10228 |  |  | Standard measuring machines/
thread measuring wire gauges |  | (0 ~ 100) mm |  | SS-CAL-10228
Thread measuring wire gauges |  | (0.17 ~ 3.2) mm | 0.7 μm | SS-CAL-10228-1
Radius gauges | 10229 |  |  | Non-Contact coordinate measuring
machines/
|  | (0.1 ~ 100) mm | 0.02 mm | SS-CAL-10229
Cylindrical ring gauges | 10230 |  |  | Standard measuring machines/
|  | (2 ~ 100) mm |  | SS-CAL-10230
Step gauges/Caliper testers | 10232 |  |  | Gauge blocks,
Electronic micrometers/
Caliper testers |  | (0 ~ 600) mm |  | SS-CAL-10232
Taper thickness gauges | 10233 |  |  | Non-Contact coordinate measuring
machines/
|  | (0 ~ 60) mm | 0.01 mm | SS-CAL-10233
Ultrasonic thickness gauges | 10234 |  |  | Ultrasonic thickness specimens/
|  | (0 ~ 100) mm | 9 μm | SS-CAL-10234
Ultrasonic/coating thickness specimens | 10235 |  |  | Electronic micrometers/
Coating thickness specimens |  | (0 ~ 2) mm | 0.4 μm | SS-CAL-10235
Ultrasonic thickness specimens |  | (0 ~ 100) mm | 2 μm | SS-CAL-10235-1
Coating thickness testers | 10236 |  |  | Gauge blocks/
|  | (0 ~ 2) mm | 3.1 μm | SS-CAL-10236
103. Angle
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Bevel protractors | 10304 |  |  | Angle Gauge Block, Gauge Block
Angle |  | ±(0 ~ 90)° | 1' | Profile Projector, Rotary Table
Angle of Accessory |  | (0 ~ 90)° | 4' | Electronic Micrometer/
SS-CAL-10304
Plate/Square/Electric levels | 10311 |  |  | Level Comparator, Gauge Block
(Plate/Square levels) |  |  |  | Cylindrical Square
Angle |  | ±(0 ~ 206)″ | 1″ | Electronic micrometers/
Squareness |  | (0 ~ 300) mm | 4.7 μm | SS-CAL-10311
Flatness |  | (0 ~ 300) mm | 3 μm
104. From
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Precision surface plates | 10407 |  |  | Electrical levels/
|  | (0 ~ 0.25) m2 | 1.8 μm | SS-CAL-10407
|  | (0.25 ~ 1) m2 | 2.4 μm
|  | (1 ~ 2) m2 | 4.2 μm
|  | (2 ~ 3) m2 | 4.4 μm
|  | (3 ~ 6) m2 | 5.0 μm
105. Complex geometry
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Contact coordinate measuring machines | 10503 |  |  | Step Gauge/
Length |  | (0 ~ 1 000) mm |  | SS-CAL-10503
Straightness |  | (0 ~ 900) mm | 2.2 μm
Non-contact coordinate measuring machines | 10504 |  |  | Standard scale/
Length |  | (0 ~ 1 000) mm |  | SS-CAL-10504
Measuring microscopes, profile projectors | 10511 |  |  | Standard scale/
Length |  | (0 ~ 500) mm |  | SS-CAL-10511
Magnification |  | (5 ~ 100) mm | 0.03 %
Angle |  | (0 ~ 360)˚ | 1.6′
Micro measuring microscopes | 10512 |  |  | Standard scale/
|  | (0 ~ 12) mm | 2.0 μm | SS-CAL-10512
Stylus Type Roughness Testers | 10517 |  |  | Roughness Specimen/
Ra |  | (0 ~ 4) μm | 0.048 μm | SS-CAL-10517
Rz |  | (0 ~ 12) μm | 0.18 μm
H |  | (0 ~ 10) μm | 0.21 μm
Roughness Standard/comparison specimens | 10519 |  |  | Roughness Tester/
(Standard specimens) |  |  |  | SS-CAL-10519
Ra |  | (0 ~ 4) μm | 0.056 μm
Rz |  | (0 ~ 12) μm | 0.20 μm
H |  | (0 ~ 10) μm | 0.27 μm
105. Complex geometry
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Thread plug gauges | 10525 |  |  | Measuring Machine 3-Wire gauge
Effective diameter |  | (1.4 ~ 100) mm | 4 μm | Form Tester, Profileprojector/
Outside diameter |  | (1.4 ~ 100) mm | 2 μm | SS-CAL-10526
Pitch |  | (0.3 ~ 5.5) mm | 3 μm
Half angle |  | (0 ~ 30)° | 3′
106. Various dimensional
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Inside/outside/gear tooth | 10601 |  |  | Caliper tersters/
calipers, caliper gauges |  |  |  | SS-CAL-10601
Inside and outside caliper |  | (0 ~ 600) mm | 0.01 mm
|  | (600 ~ 1 000) mm | 0.02 mm
Caliper gauges |  | (0 ~ 600) mm | 0.01 mm
Cylinder/bore gauges | 10603 |  |  | Dial gauge testers/
|  | (0 ~ 2) mm | 0.9 μm | SS-CAL-10603
Depth gauges, depth micrometers | 10604 |  |  | Gauge blocks/
Depth gauges |  | (0 ~ 600) mm | 0.01 mm | SS-CAL-10604
Depth micrometers |  | (0 ~ 300) mm
Dial/digital gauges | 10605 |  |  | Gauge blocks/
|  | (0 ~ 100) mm |  | SS-CAL-10605
Micro indicators, test indicators | 10609 |  |  | Dial gauge testers/
|  | (0 ~ 5) mm | 0.5 μm | SS-CAL-10609
Micrometer head | 10610 |  |  | Gauge blocks/
|  | (0 ~ 25) mm | 2 μm | SS-CAL-10610
3-point micrometers | 10611 |  |  | Cylindrical ring gauges/
|  | (2 ~ 100) mm | 1.5 μm | SS-CAL-10611
Inside micrometers | 10612 |  |  | Gauge blocks/
Inside micrometers |  | (5 ~ 300) mm |  | SS-CAL-10612
Bar type micrometers |  | (50 ~ 500) mm
Outside micrometers | 10613 |  |  | Gauge blocks,
Outside micrometers |  | (0 ~ 600) mm |  | Cylindrical plug gauges/
V-anvil micrometers |  | (1 ~ 25) mm | 1.4 μm | SS-CAL-10613
Test sieve | 10617 |  |  | Non-contact coordinate measuring machines /
Wire Rod Diameter |  | (0 ~ 10) mm | 4.4 μm | SS-CAL-10617
Opening size |  | (0 ~ 130) mm | 6.4 μm
201. Mass
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electric balances | 20109 |  |  | Weight/
|  | (0 ~ 2) g | 30 μg | SS-CAL-20109
|  | (2 ~ 20) g | 58 μg
|  | (20 ~ 300) g | 0.25 mg
|  | (300 ~ 500) g | 0.53 mg
|  | (500 ~ 2 000) g | 1.8 mg
201. Mass
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electric balances | 20109 |  |  | Weight/
|  | (2 ~ 10) kg | 10 mg | SS-CAL-20109
|  | (10 ~ 30) kg | 22 mg
|  | (30 ~ 60) kg | 2.0 g
|  | (60 ~ 150) kg | 2.2 g
|  | (150 ~ 300) kg | 4.6 g
|  | (300 ~ 500) kg | 44 g
|  | (500 ~ 1 000) kg | 0.11 kg
Spring scale balances | 20113 |  |  | Weights/
|  | (0 ~ 1) kg | 1.6 g | SS-CAL-20113
|  | (1 ~ 10) kg | 16 g
|  | (10 ~ 50) kg | 0.07 kg
|  | (50 ~ 100) kg | 0.16 kg
Weights | 20116 |  |  | Weight, Electric balances/
(F1 or less) |  | 1 mg | 2.4 μg | SS-CAL-20116
|  | 2 mg | 2.4 μg
|  | 5 mg | 2.4 μg
|  | 10 mg | 3.1 μg
|  | 20 mg | 3.8 μg
|  | 50 mg | 4.7 μg
|  | 100 mg | 5.4 μg
|  | 200 mg | 7.0 μg
|  | 500 mg | 8.5 μg
|  | 1 g | 9.8 μg
|  | 2 g | 14 μg
|  | 5 g | 21 μg
|  | 10 g | 24 μg
|  | 20 g | 29 μg
|  | 50 g | 38 μg
|  | 100 g | 58 μg
|  | 200 g | 0.13 mg
|  | 500 g | 0.38 mg
|  | 1 kg | 0.59 mg
|  | 2 kg | 1.1 mg
|  | 5 kg | 2.7 mg
|  | 10 kg | 5.4 mg
|  | 20 kg | 11 mg
202. Force
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Tension/compression testing machines | 20203 |  |  | Electrical force measuring devices/
Tension |  | (1 ~ 10) N | 1.9 × 10-3 | SS-CAL-20203
|  | (10 ~ 100) N | 1.9 × 10-3
|  | (100 ~ 200) N | 1.9 × 10-3
|  | (200 ~ 500) N | 2.0 × 10-3
|  | (0.5 ~ 1) kN | 1.8 × 10-3
|  | (1 ~ 2) kN | 1.7 × 10-3
|  | (2 ~ 5) kN | 1.9 × 10-3
|  | (5 ~ 10) kN | 2.0 × 10-3
|  | (10 ~ 20) kN | 2.0 × 10-3
202. Force
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Tension/compression testing machines | 20203 |  |  | Electrical force measuring devices/
Compression |  | (1 ~ 10) N | 1.9 × 10-3 | SS-CAL-20203
|  | (10 ~ 100) N | 1.9 × 10-3
|  | (100 ~ 200) N | 1.7 × 10-3
|  | (200 ~ 500) N | 1.6 × 10-3
|  | (0.5 ~ 1) kN | 1.7 × 10-3
|  | (1 ~ 2) kN | 1.8 × 10-3
|  | (2 ~ 5) kN | 1.6 × 10-3
|  | (5 ~ 10) kN | 1.7 × 10-3
|  | (10 ~ 20) kN | 2.2 × 10-3
|  | (20 ~ 50) kN | 1.9 × 10-3
|  | (50 ~ 100) kN | 1.9 × 10-3
|  | (100 ~ 200) kN | 1.9 × 10-3
|  | (200 ~ 500) kN | 2.0 × 10-3
|  | (500 ~ 1 000) kN | 2.2 × 10-3
|  | (1 000 ~ 2 000) kN | 2.2 × 10-3
Push-pull gauges | 20204 |  |  | Weight/
|  | (1 ~ 1 000) N | 1.3 × 10-3 | SS-CAL-20204
203. Torque
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Torque wrenches/drivers | 20303 |  |  | Torque measuring devices/
|  | (0.06 ~ 0.6) N·m | 2.8 × 10-2 | SS-CAL-20303
|  | (0.6 ~ 1.0) N·m | 1.1 × 10-2
|  | (1 ~ 2.5) N·m | 7.7 × 10-3
|  | (2.5 ~ 5) N·m | 7.3 × 10-3
|  | (5 ~ 10) N·m | 6.5 × 10-3
|  | (10 ~ 25) N·m | 5.4 × 10-3
|  | (25 ~ 50) N·m | 5.2 × 10-3
|  | (50 ~ 100) N·m | 8.9 × 10-3
|  | (100 ~ 250) N·m | 5.4 × 10-3
|  | (250 ~ 500) N·m | 5.1 × 10-3
|  | (500 ~ 1 000) N·m | 5.7 × 10-3
204. Pressure
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Manometers | 20402 |  |  | Pressure Controller,
|  | (0 ~ 10) k㎩ | 6.6 × 10-3 | Pressure Balance/
|  | (10 ~ 200) k㎩ | 6.5 × 10-3 | SS-CAL-20402
Pneumatic Pressuce Balance | 20403 |  |  | Pressure Balance/
|  | (4 ~ 700) k㎩ | 8.6 × 10-5 | SS-CAL-20403
|  | (700 ~ 3 500) k㎩ | 6.0 × 10-5
|  | (3 500 ~ 7 000) k㎩ | 8.5 × 10-5
Hydraulic Pressure Balance | 20404 |  |  | Pressure Balance/
|  | (0.1 ~ 100) M㎩ | 6.9 × 10-5 | SS-CAL-20404
|  | (100 ~ 200) M㎩ | 7.4 × 10-5
Absolute Pressure Gauges | 20406 |  |  | Pressure Balance/
|  | (4 ~ 200) k㎩ abs. | 8.6 × 10-5 | SS-CAL-20406
|  | (200 ~ 7 000) k㎩ abs. | 8.2 × 10-5
Compound Pressure Gauges | 20408 |  |  | Pressure Balance/
|  | -95 k㎩ ~ 7 M㎩ | 2.1 × 10-4 | SS-CAL-20408
Differential Pressure Gauges | 20409 |  |  | Digital Pressure Gauge,
|  | (0 ~ 10) k㎩ | 1.4 × 10-3 | Pressure Balance/
|  | 10 k㎩ ~ 7 M㎩ | 8.7 × 10-5 | SS-CAL-20409
Gauge Pressure Gauges | 20411 |  |  | Pressure Balance/
|  | (0 ~ 700) k㎩ | 8.5 × 10-5 | SS-CAL-20411
|  | (700 ~ 7 000) k㎩ | 8.2 × 10-5
|  | 7 000 ㎪ ~ 100 M㎩ | 7.1 × 10-5
|  | (100 ~ 200) M㎩ | 8.8 × 10-5
Pressure | 20412 |  |  | Pressure Balance/
Transducers/Transmitter |  | (0 ~ 700) k㎩ | 1.3 × 10-4 | SS-CAL-20412
|  | (700 ~ 7 000) k㎩ | 1.2 × 10-4
|  | 7 000 ㎪ ~ 100 M㎩ | 1.1 × 10-4
|  | (100 ~ 200) M㎩ | 1.6 × 10-4
|  | (-95 ~ 0) kPa | 2.0 × 10-4
|  | (4 ~ 200) kPa abs. | 3.0 × 10-4
|  | (200 ~ 7 000) kPa abs. | 1.2 × 10-4
Dial Type Vaccum Gauges | 20413 |  |  | Pressure Balance/
|  | (-95 ~ 0) k㎩ | 1.0 × 10-4 | SS-CAL-20413
206. Volume
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Volumetric glasswares | 20601 |  |  | Precision balances, Weights, Ultrapure water
|  | (0.1 ~ 1) ml | 1.2 μl | SS-CAL-20601
|  | (1 ~ 2) ml | 2.4 μl
|  | (2 ~ 5) ml | 5.8 μl
|  | (5 ~ 25) ml | 12 μl
|  | (25 ~ 50) ml | 39 μl
|  | (50 ~ 100) ml | 40 μl
|  | (100 ~ 250) ml | 53 μl
|  | (250 ~ 500) ml | 89 μl
|  | (500 ~ 1 000) ml | 0.15 ml
|  | (1 000 ~ 2 000) ml | 0.26 ml
|  | (2 000 ~ 5 000) ml | 0.58 ml
Piston type volume meters | 20606 |  |  | Precision balances, Weights, Ultrapure water
|  | (0.1 ~ 2) μl | 12 nl | SS-CAL-20606
|  | (2 ~ 5) μl | 16 nl
|  | (5 ~ 10) μl | 26 nl
|  | (10 ~ 20) μl | 42 nl
|  | (20 ~ 50) μl | 86 nl
|  | (0.05 ~ 0.1) ml | 0.14 μl
|  | (0.1 ~ 0.2) ml | 0.25 μl
|  | (0.2 ~ 0.5) ml | 0.75 μl
|  | (0.5 ~ 1) ml | 1.4 μl
|  | (1 ~ 2) ml | 2.5 μl
|  | (2 ~ 5) ml | 5.7 μl
|  | (5 ~ 10) ml | 11 μl
|  | (10 ~ 20) ml | 23 μl
301. Time/frequency
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Frequency standards | 30102 |  |  | GPS receiver. Counters/
Timebase Frequency |  | (1 ~ 10) M㎐ | 6.0 × 10-12 | SS-CAL-30102
General frequency sources | 30103 |  |  | GPS receiver. Counters/
Timebase Frequency |  | (0.1 ~ 10) M㎐ | 6.0 × 10-12 | SS-CAL-30103
Frequency meters/counters | 30104 |  |  | GPS receiver. Counters
Timebase Frequency |  | (1 ~ 10) M㎐ | 6.0 × 10-12 | General frequency sources /
SS-CAL-30104
Input Frequency |  | (1 ~ 100) k㎐ | 1.4 × 10-10
|  | 100 k㎐ ~ 100 M㎐ | 6.6 × 10-12
|  | 100 M㎐ ~ 26.5 G㎐ | 6.2 × 10-11
Time interval meters/Stop watches/Timers | 30106 |  |  | Stopwatch calibration/
Relative time |  | Day | 2.0 × 10-7 | SS-CAL-30106
|  | Month | 6.0 × 10-8
Time |  | (1 ~ 10 000) s | 5.8 × 10-6
302. Velocity & revolution
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Standard RPM generators | 30201 |  |  | GPS receiver. Function
Revolution Velocity Measurement |  | (1 ~ 4 000) min-1 | 5.8 × 10-2 ~min-1 | Generator /
SS-CAL-30201
Revolution Velocity Measurement |  | 100 min-1 | 0.01 min-1
(Centrifuge) |  | (100 ~ 400) min-1 | 0.02 min-1
|  | (400 ~ 700) min-1 | 0.03 min-1
|  | (700 ~ 1 000) min-1 | 0.05 min-1
|  | (1 000 ~ 3 000) min-1 | 0.14 min-1
|  | (3 000 ~ 7 000) min-1 | 0.32 min-1
|  | (7 000 ~ 11 000) min-1 | 0.5 min-1
|  | (11 000 ~ 13 000) min-1 | 0.6 min-1
|  | (13 000 ~ 15 000) min-1 | 0.7 min-1
|  | (15 000 ~ 18 000) min-1 | 0.8 min-1
|  | (18 000 ~ 20 000) min-1 | 0.9 min-1
|  | (20 000 ~ 30 000) min-1 | 3.6 min-1
|  | (30 000 ~ 40 000) min-1 | 4.7 min-1
|  | (40 000 ~ 50 000) min-1 | 5.9 min-1
|  | (50 000 ~ 60 000) min-1 | 7.0 min-1
|  | (60 000 ~ 70 000) min-1 | 8.1 min-1
|  | (70 000 ~ 80 000) min-1 | 9.3 min-1
|  | (80 000 ~ 90 000) min-1 | 10 min-1
|  | (90 000 ~ 99 000) min-1 | 11 min-1
Contact type tachometers | 30202 |  |  | GPS Receiver, Tachometer Cal
Revolution Velocity Measurement |  | (10 ~ 4 000) min-1 | 0.12 min-1 | System/
SS-CAL-30202
Photo tachometers/stroboscopes | 30203 |  |  | GPS Receiver, Photo Signal
Revolution Velocity Measurement |  | (6 ~ 90) min-1 | 0.01 min-1 | Generator, Detector
(Photo tachometers) |  | (90 ~ 900) min-1 | 0.1 min-1 | SS-CAL-30203
|  | (900 ~ 90 000) min-1 | 1 min-1
Revolution Velocity Measurement |  | (6 ~ 9 000) min-1 | 0.01 min-1
(Stroboscopes) |  | (9 000 ~ 90 000) min-1 | 0.1 min-1
|  | (90 000 ~ 200 000) min-1 | 1 min-1
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC ammeters | 40101 | (±) |  | Calibrator/
DC Current |  | 1 nA | 1.4 × 10-2 | SS-CAL-40101
|  | (1 ~ 100) nA | 8.4 × 10-3
|  | 100 ㎁ ~ 1 μA | 2.4 × 10-3
|  | (1 ~ 10) μA | 2.0 × 10-4
|  | (10 ~ 100) μA | 4.0 × 10-5
|  | 100 μA ~ 1 mA | 3.6 × 10-5
|  | (1 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 6.6 × 10-5
|  | (1 ~ 10) A | 4.2 × 10-4
|  | (10 ~ 20) A | 1.5 × 10-4
|  | (20 ~ 30) A | 3.3 × 10-4
|  | (30 ~ 50) A | 2.4 × 10-4
|  | (50 ~ 100) A | 1.6 × 10-4
Transconductance amplifiers | 40102 | (±) |  | AC-DC Active Current Shunt/
DC Current |  | 100 μA | 3.8 × 10-6 | SS-CAL-40102
|  | 100 μA ~ 100 mA | 3.8 × 10-6
|  | 100 mA ~ 1 A | 4.2 × 10-6
|  | (1 ~ 10) A | 6.2 × 10-6
|  | (10 ~ 20) A | 2.5 × 10-5
|  | (20 ~ 50) A | 1.2 × 10-5
|  | (50 ~ 100) A | 3.0 × 10-5
AC Current |  | (40 ~ 60) ㎐
|  | 100 μA | 1.1 × 10-4
|  | 100 μA ~ 1 mA | 6.2 × 10-5
|  | (1 ~ 10) mA | 1.0 × 10-4
|  | (10 ~ 100) mA | 4.4 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.1 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.8 × 10-5
(60 ~ 400) ㎐
|  | 100 μA | 1.6 × 10-4
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 6.6 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.4 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
400 ㎐ ~ 1 ㎑
|  | 100 μA | 8.8 × 10-5
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 6.6 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Transconductance amplifiers | 40102 |  |  | AC-DC Active Current Shunt/
AC Current |  | (1 ~ 10) ㎑ |  | SS-CAL-40102
|  | 100 μA | 1.2 × 10-4
|  | 100 μA ~ 1 mA | 5.8 × 10-5
|  | (1 ~ 10) mA | 7.4 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 6.9 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
DC voltage/current calibrators | 40103 | (±) |  | Standard Resistor, Calibrator
DC Voltage |  | 1 mV | 3.8 × 10-4 | Reference Multimeter, Shunt,
|  | (1 ~ 10) mV | 3.8 × 10-5 | AC Measurement STD/
|  | (10 ~ 100) mV | 3.6 × 10-6 | SS-CAL-40103
|  | 100 mV ~ 1 V | 6.6 × 10-7
|  | (1 ~ 10) V | 4.6 × 10-7
|  | (10 ~ 100) V | 5.6 × 10-7
|  | (100 ~ 1 000) V | 8.0 × 10-7
DC Current |  | (±)
|  | 1 nA | 1.1 × 10-2
|  | (1 ~ 100) nA | 7.0 × 10-3
|  | 100 nA ~ 1 μA | 5.8 × 10-5
|  | (1 ~ 10) μA | 7.0 × 10-6
|  | 10 μA ~ 100 mA | 3.8 × 10-6
|  | 100 mA ~ 1 A | 4.2 × 10-6
|  | (1 ~ 10) A | 6.2 × 10-6
|  | (10 ~ 20) A | 2.5 × 10-5
|  | (20 ~ 50) A | 1.2 × 10-5
|  | (50 ~ 100) A | 3.0 × 10-5
Electricaltemperaturecalibrators | 40104 |  |  | Digital Multimeter,Calibraror/
DC VOLTAGE(SOURCE) |  | (±) |  | SS-CAL-40104
|  | 10 mV | 6.0 × 10-5
|  | (10 ~ 100) mV | 8.0 × 10-6
|  | 100 mV ~ 100 V | 8.0 × 10-6
DC CURRENT(SOURCE) |  | (±)
|  | 1 mA | 2.0 × 10-4
|  | (1 ~ 10) mA | 2.0 × 10-5
|  | (10 ~ 100) mA | 4.8 × 10-5
RESISTANCE(SOURCE) |  | 1 Ω | 6.0 × 10-5
|  | 1 ~ 10 Ω | 6.0 × 10-5
|  | (10 ~ 100) Ω | 1.2 × 10-5
|  | 100 Ω ~ 1 kΩ | 1.2 × 10-5
|  | 1 kΩ ~ 10 kΩ | 2.0 × 10-5
|  | (10 ~ 100) kΩ | 1.0 × 10-5
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electricaltemperaturecalibrators | 40104 |  |  | Digital Multimeter,Calibraror/
DC VOLTAGE(MEASURE) |  | (±) |  | SS-CAL-40104
|  | 10 mV | 1.0 × 10-4
|  | (10 ~ 100) mV | 1.6 × 10-5
|  | 100 mV ~ 1 V | 8.0 × 10-6
|  | (1 ~ 10) V | 5.4 × 10-6
|  | (10 ~ 100) V | 1.0 × 10-5
|  | (100 ~ 300) V | 1.1 × 10-5
DC CURRENT(MEASURE) |  | (±)
|  | 1 mA | 8.0 × 10-5
|  | (1 ~ 10) mA | 5.0 × 10-5
|  | (10 ~ 100) mA | 6.8 × 10-5
RESISTANCE(MEASURE) |  | 1 Ω | 2.0 × 10-4
|  | (1 ~ 10) Ω | 8.0 × 10-5
|  | (10 ~ 100) Ω | 2.0 × 10-5
|  | 100 Ω ~ 100 kΩ | 4.0 × 10-5
DC current shunts | 40105 |  |  | Trans Conductance Amplifier/
Resistance |  | 50 μΩ | 1.6 × 10-4 | SS-CAL-40105
|  | 50 μΩ ~ 1 mΩ | 2.0 × 10-4
|  | 1 mΩ ~ 10 mΩ | 1.4 × 10-4
|  | 10 mΩ ~ 100 mΩ | 1.1 × 10-4
|  | 100 mΩ ~ 1 Ω | 6.0 × 10-5
|  | 1 Ω ~ 10 Ω | 6.0 × 10-5
|  | 10 Ω ~ 100 Ω | 5.4 × 10-5
|  | 100 Ω ~ 1 kΩ | 1.2 × 10-4
|  | 1 kΩ ~ 10 kΩ | 7.4 × 10-4
DC power supplies | 40108 | (±) |  | DC Electronics Load/
DC Voltage |  | 10 mV | 6.0 × 10-5 | SS-CAL-40108
|  | (10 ~ 100) mV | 8.0 × 10-6
|  | 100 mV ~ 10 V | 6.8 × 10-6
|  | (10 ~ 100) V | 8.0 × 10-6
|  | (100 ~ 300) V | 6.7 × 10-6
|  | (300 ~ 500) V | 1.2 × 10-5
|  | (500 ~ 1 000) V | 8.0 × 10-6
DC Current |  | 1 mA | 6.0 × 10-4
|  | (1 ~ 10) mA | 8.0 × 10-5
|  | (10 ~ 200) mA | 7.0 × 10-5
|  | 200 mA ~ 2 A | 1.2 × 10-4
|  | (2 ~ 20) A | 4.8 × 10-4
|  | (20 ~ 100) A | 2.0 × 10-4
|  | (100 ~ 200) A | 4.0 × 10-4
|  | (200 ~ 300) A | 4.0 × 10-4
|  | (300 ~ 600) A | 4.0 × 10-4
|  | (600 ~ 1 000) A | 3.8 × 10-4
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC voltmeters | 40112 |  |  | Calibrator/
DC Voltage |  | (±) |  | SS-CAL-40112
|  | 1 mV | 6.8 × 10-4
|  | (1 ~ 10) mV | 7.0 × 10-5
|  | (10 ~ 100) mV | 8.2 × 10-6
|  | 100 mV ~ 1 V | 3.4 × 10-6
|  | (1 ~ 10) V | 2.0 × 10-6
|  | (10 ~ 100) V | 2.8 × 10-6
|  | (100 ~ 1 000) V | 3.2 × 10-6
Static/Ionicvoltmeters | 40113 |  |  | DC Power Supply/
DC Voltage |  | (±) |  | SS-CAL-40113
|  | 0.1 kV | 6.0 × 10-3
|  | (0.1 ~ 50) kV | 3.0 × 10-3
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Capacitance bridges/indicators | 40201 |  |  | Standard Capacitance Set/
Frequency |  | 60 ㎐ ~ 100 M㎐ | 6.2 × 10-8 | SS-CAL-40201
Capacitance |  | 1 pF
|  | 1 k㎐ ~ 1 M㎐ | 3.6 × 10-4
|  | (1 ~ 4) M㎐ | 7.2 × 10-4
|  | 5 M㎐ | 9.4 × 10-4
|  | 10 M㎐ | 2.4 × 10-3
|  | 13 M㎐ | 3.7 × 10-3
10 pF
|  | 1 k㎐ ~ 10 M㎐ | 3.6 × 10-4
|  | 13 M㎐ | 3.8 × 10-4
100 pF
|  | 1 k㎐ ~ 4 M㎐ | 3.6 × 10-4
|  | 4 M㎐ ~ 5 M㎐ | 3.8 × 10-4
|  | 5 M㎐ ~ 10 M㎐ | 4.8 × 10-4
|  | 13 M㎐ | 6.0 × 10-4
1 000 pF
|  | 1 k㎐ ~ 1 M㎐ | 3.6 × 10-4
|  | (1 ~ 4) M㎐ | 5.6 × 10-4
|  | 5 M㎐ | 7.2 × 10-4
|  | 10 M㎐ | 1.9 × 10-3
|  | 13 M㎐ | 2.9 × 10-3
10 ㎋
|  | 120 ㎐ ~ 100 k㎐ | 7.8 × 10-5
100 ㎋
|  | 120 ㎐ ~ 100 k㎐ | 7.8 × 10-5
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Capacitance bridges/indicators | 40201 |  |  | Standard Capacitance Set/
Capacitance |  | 1 μF |  | SS-CAL-40201
|  | 120 ㎐ | 8.4 × 10-5
|  | 1 k㎐ ~ 10 k㎐ | 7.8 × 10-5
|  | 100 k㎐ | 9.7 × 10-5
10 μF ~ 10 mF
|  | 100 ㎐ | 6.0 × 10-4
|  | 120 ㎐ | 6.0 × 10-4
10 μF
|  | 1 k㎐ | 8.0 × 10-4
100 μF
|  | 1 k㎐ | 8.0 × 10-4
1 mF
|  | 1 k㎐ | 1.0 × 10-3
10 mF
|  | 1 k㎐ | 2.4 × 10-3
Decade capacitors | 40202 |  |  | Standard Capacitance Set/
Capacitance |  | at 1 k㎐ |  | SS-CAL-40202
|  | 1 pF ~ 1.2 μF | 3.5 × 10-3
Earth testers | 40205 |  |  | Decade Resistor/
|  | (50 ㎐, 60 ㎐) |  | SS-CAL-40205
AC Voltage |  | (1 ~ 10) V | 6.4 × 10-4
|  | (10 ~ 30) V | 2.6 × 10-4
|  | (30 ~ 50) V | 1.6 × 10-4
|  | (50 ~ 70) V | 1.3 × 10-4
|  | (70 ~ 100) V | 6.4 × 10-4
|  | (100 ~ 500) V | 2.3 × 10-4
|  | (500 ~ 600) V | 1.4 × 10-4
DC Voltage |  | (1 ~ 10) V | 6.4 × 10-4
|  | (10 ~ 30) V | 2.1 × 10-4
|  | (30 ~ 50) V | 1.3 × 10-4
|  | (50 ~ 70) V | 9.1 × 10-4
|  | (70 ~ 100) V | 6.4 × 10-4
|  | (100 ~ 600) V | 2.1 × 10-4
|  | (600 ~ 1 000) V | 9.1 × 10-4
Resistance |  | 1 mΩ ~ 1 Ω | 2.6 × 10-3
|  | (1 ~ 10) Ω | 6.8 × 10-4
|  | (10 ~ 100) Ω | 6.6 × 10-4
|  | 100 Ω ~ 1 kΩ | 6.4 × 10-4
|  | (1 ~ 100) kΩ | 6.0 × 10-4
(50 ㎐, 60 ㎐)
AC Current out |  | (1 ~ 3) A | 6.7 × 10-4
|  | (3 ~ 10) A | 6.6 × 10-4
|  | (10 ~ 20) A | 6.2 × 10-4
|  | (20 ~ 30) A | 6.7 × 10-4
|  | (30 ~ 60) A | 6.3 × 10-4
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Earth testers | 40205 |  |  | Decade Resistor/
Timer |  | (1 ~ 100) s | 8.2 × 10-6 | SS-CAL-40205
|  | (100 ~ 1 000) s | 5.8 × 10-5
Inductance bridges/indicators | 40206 |  |  | Standard Inductor/
Frequency |  | 60 ㎐ ~ 100 M㎐ | 6.2 × 10-8 | SS-CAL-40206
Inductance |  | at 1 k㎐
|  | 10 μH | 1.2 × 10-3
|  | (10 ~ 100) μH | 1.9 × 10-4
|  | (0.1 ~ 100) mH | 1.3 × 10-4
|  | (1 ~ 10) H | 1.6 × 10-4
Inductors | 40208 |  |  | Standard Inductor/
Standard Inductance |  | at 1 k㎐ |  | SS-CAL-40208
|  | 10 μH | 1.2 × 10-3
|  | 100 μH | 2.8 × 10-4
|  | 1 mH | 2.4 × 10-4
|  | 10 mH | 2.4 × 10-4
|  | 100 mH | 2.4 × 10-4
|  | 1 H | 2.4 × 10-4
|  | 10 H | 2.6 × 10-4
Decade Inductance |  | at 1 k㎐
|  | 100 μH ~ 10 H | 3.5 × 10-3
Insulation testers | 40210 | (±) |  | High Resistance Decade/
DC Voltage |  | (1 ~ 10) V | 6.4 × 10-4 | SS-CAL-40210
|  | (10 ~ 30) V | 2.1 × 10-4
|  | (30 ~ 100) V | 6.4 × 10-5
|  | (100 ~ 300) V | 2.3 × 10-5
|  | (300 ~ 500) V | 1.6 × 10-5
|  | (500 ~ 700) V | 1.2 × 10-5
|  | (700 ~ 1 000) V | 6.4 × 10-5
AC Voltage |  | (50 ㎐, 60 ㎐)
|  | (1 ~ 10) V | 6.4 × 10-4
|  | (10 ~ 30) V | 2.6 × 10-4
|  | (30 ~ 100) V | 1.0 × 10-4
|  | (100 ~ 300) V | 1.2 × 10-4
|  | (300 ~ 500) V | 1.0 × 10-4
|  | (500 ~ 700) V | 9.4 × 10-5
|  | (700 ~ 1 000) V | 1.1 × 10-4
Insulation Voltage |  | (1 ~ 50) V | 1.3 × 10-3
|  | (50 ~ 100) V | 6.4 × 10-4
|  | (100 ~ 250) V | 2.1 × 10-4
|  | (250 ~ 500) V | 1.3 × 10-4
|  | (500 ~ 700) V | 9.1 × 10-5
|  | 700 V ~ 1 ㎸ | 6.4 × 10-5
|  | (1 ~ 5) ㎸ | 7.6 × 10-3
|  | (5 ~ 10) ㎸ | 6.8 × 10-3
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Insulation testers | 40210 |  |  | High Resistance Decade/
Insulation Resistance |  | 1 kΩ | 6.6 × 10-4 | SS-CAL-40210
|  | (1 ~ 10) kΩ | 7.2 × 10-5
|  | (10 ~ 100) kΩ | 3.2 × 10-5
|  | 100 kΩ ~ 1 MΩ | 6.4 × 10-4
|  | (1 ~ 10) MΩ | 8.6 × 10-5
|  | (10 ~ 100) MΩ | 1.7 × 10-4
|  | 100 MΩ ~ 1 GΩ | 8.8 × 10-4
|  | (1 ~ 10) GΩ | 1.2 × 10-3
|  | (10 ~ 100) GΩ | 1.7 × 10-3
|  | 100 GΩ ~ 1 TΩ | 2.4 × 10-3
Timer |  | (1 ~ 100) s | 5.8 × 10-6
|  | (100 ~ 1 000) s | 8.2 × 10-6
|  | (1 000 ~ 10 000) s | 5.8 × 10-4
Resistance bridges & | 40213 |  |  | Standard Resistance Set/
similar instruments |  | (1 ~ 100) mΩ | 6.2 × 10-3 | SS-CAL-40213
Resistance(Rheostat Arm) |  | 100 mΩ ~ 1 Ω | 6.2 × 10-4
|  | (1 ~ 10) Ω | 6.4 × 10-5
|  | (10 ~ 100) Ω | 2.3 × 10-5
|  | 100 Ω ~ 1 kΩ | 2.2 × 10-5
|  | (1 ~ 10) kΩ | 6.2 × 10-5
|  | (10 ~ 100) kΩ | 2.3 × 10-5
|  | 100 kΩ ~ 1 MΩ | 2.7 × 10-5
|  | (1 ~ 10) MΩ | 9.4 × 10-5
|  | (10 ~ 100) MΩ | 6.5 × 10-4
Resistance(Ratio Arm) |  | (1 ~ 100) mΩ | 6.2 × 10-5
|  | 100 mΩ ~ 10 Ω | 9.8 × 10-6
|  | (10 ~ 100) Ω | 7.8 × 10-6
|  | 100 Ω ~ 1 kΩ | 7.4 × 10-6
|  | (1 ~ 10) kΩ | 5.0 × 10-6
|  | (10 ~ 100) kΩ | 7.4 × 10-6
|  | 100 kΩ ~ 1 MΩ | 9.6 × 10-6
|  | (1 ~ 10) MΩ | 1.2 × 10-5
|  | (10 ~ 100) MΩ | 2.5 × 10-5
Resistance meters | 40214 |  |  | Standard Resistance Set/
DC Resistance |  | 50 μΩ | 2.3 × 10-3 | SS-CAL-40214
|  | (50 ~ 100) μΩ | 2.4 × 10-3
|  | 100 μΩ ~ 1 mΩ | 6.2 × 10-5
|  | (1 ~ 10) mΩ | 3.2 × 10-5
|  | (10 ~ 100) mΩ | 2.6 × 10-5
|  | 100 mΩ ~ 1 Ω | 9.8 × 10-6
|  | (1 ~ 10) Ω | 9.8 × 10-6
|  | (10 ~ 100) Ω | 7.8 × 10-6
|  | 100 Ω ~ 1 kΩ | 7.4 × 10-6
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Resistance meters | 40214 |  |  | Standard Resistance Set/
DC Resistance |  | (1 ~ 10) kΩ | 5.0 × 10-6 | SS-CAL-40214
|  | (10 ~ 100) kΩ | 7.4 × 10-6
|  | 100 kΩ ~ 1 MΩ | 9.6 × 10-6
|  | (1 ~ 10) MΩ | 1.3 × 10-5
|  | (10 ~ 100) MΩ | 2.4 × 10-5
|  | 100 MΩ ~ 1 GΩ | 2.4 × 10-4
|  | (1 ~ 10) GΩ | 1.2 × 10-3
|  | (10 ~ 100) GΩ | 1.7 × 10-3
|  | 100 GΩ ~ 1 TΩ | 2.4 × 10-3
Frequency |  | 1 k㎐ | 6.2 × 10-7
AC Voltage |  | (10 ~ 100) mV | 1.4 × 10-4
|  | 100 mV ~ 10 V | 9.6 × 10-5
AC Resistance |  | at 1 k㎐
|  | 1 mΩ | 5.0 × 10-3
|  | (1 ~ 10) mΩ | 5.2 × 10-4
|  | (10 ~ 100) mΩ | 3.3 × 10-4
|  | 100 mΩ ~ 10 kΩ | 1.3 × 10-4
|  | 10 kΩ ~ 100 kΩ | 1.4 × 10-4
Resistors | 40215 |  |  | Standard Resistance Set/
Decade Resistance |  | (1 ~ 10) mΩ | 5.8 × 10-3 | SS-CAL-40215
|  | (10 ~ 100) mΩ | 6.2 × 10-3
|  | 100 mΩ ~ 1 Ω | 6.2 × 10-4
|  | (1 ~ 10) Ω | 6.4 × 10-5
|  | (10 ~ 100) Ω | 2.3 × 10-5
|  | 100 Ω ~ 1 kΩ | 2.2 × 10-5
|  | (1 ~ 10) kΩ | 6.2 × 10-5
|  | (10 ~ 100) kΩ | 2.3 × 10-5
|  | 100 kΩ ~ 1 MΩ | 2.7 × 10-5
|  | (1 ~ 10) MΩ | 9.4 × 10-5
|  | (10 ~ 100) MΩ | 1.9 × 10-4
|  | 100 MΩ ~ 1 GΩ | 1.9 × 10-4
|  | (1 ~ 10) GΩ | 8.3 × 10-4
|  | (10 ~ 100) GΩ | 9.5 × 10-4
|  | 100 GΩ ~ 1 TΩ | 6.2 × 10-4
|  | (1 ~ 10) TΩ | 9.6 × 10-4
AC Resistance |  | (40 ㎐ ~ 1 k㎐)
|  | 1 mΩ | 6.2 × 10-4
|  | 10 mΩ | 5.8 × 10-4
|  | 100 mΩ | 3.6 × 10-4
|  | 1 Ω | 2.0 × 10-4
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Resistors | 40215 |  |  | Standard Resistance Set/
AC Resistance |  | 10 Ω(1 k㎐) | 2.0 × 10-4 | SS-CAL-40215
|  | 1 M㎐ | 3.8 × 10-4
|  | 2 M㎐ | 5.6 × 10-4
|  | 100 Ω(1 k㎐) | 2.0 × 10-4
|  | 1 M㎐ | 3.8 × 10-4
|  | 2 M㎐ | 4.6 × 10-4
|  | 1 kΩ (1 k㎐) | 2.6 × 10-4
|  | 100 k㎐ | 3.8 × 10-4
|  | 1 M㎐ | 3.8 × 10-4
|  | 2 M㎐ | 3.8 × 10-4
|  | 10 kΩ (1 k㎐) | 1.1 × 10-3
|  | 100 k㎐ | 3.0 × 10-4
|  | 1 M㎐ | 3.8 × 10-4
|  | 100 kΩ (1 k㎐) | 9.8 × 10-3
|  | 100 k㎐ | 3.8 × 10-4
|  | 1 M㎐ | 3.8 × 10-4
Resistance |  | 1 Ω | 9.9 × 10-6
|  | (1 ~ 10) Ω | 9.8 × 10-6
|  | (10 ~ 100) Ω | 7.8 × 10-6
|  | 100 Ω ~ 1 kΩ | 7.4 × 10-6
|  | (1 ~ 10) kΩ | 5.0 × 10-6
|  | (10 ~ 100) kΩ | 7.4 × 10-6
|  | 100 kΩ ~ 1 MΩ | 9.6 × 10-6
|  | (1 ~ 10) MΩ | 1.2 × 10-5
|  | (10 ~ 100) MΩ | 2.4 × 10-5
|  | 100 MΩ ~ 1 GΩ | 2.4 × 10-4
Impedance bridges/LCR meters | 40217 |  |  | Standard Capacitor Set,
Frequency |  | 60 ㎐ ~ 100 M㎐ | 6.2 × 10-8 | Standard Resistor Set,
Standard Inductor/
AC Voltage |  | 10 mV | 4.2 × 10-4 | SS-CAL-40217
|  | (10 ~ 100) mV | 1.4 × 10-4
|  | 100 mV ~ 10 V | 9.6 × 10-5
|  | (10 ~ 100) V | 9.8 × 10-5
Capacitance |  | 1 pF
|  | 1 k㎐ ~ 1 M㎐ | 3.6 × 10-4
|  | (1 ~ 4) M㎐ | 7.2 × 10-4
|  | 5 M㎐ | 9.4 × 10-4
|  | 10 M㎐ | 2.4 × 10-3
|  | 13 M㎐ | 3.7 × 10-3
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Impedance bridges/LCR meters | 40217 |  |  | Standard Capacitor Set,
Capacitance |  | 10 pF |  | Standard Resistor Set,
|  | 1 k㎐ ~ 5 M㎐ | 3.6 × 10-4 | Standard Inductor/
|  | 10 M㎐ | 3.6 × 10-4 | SS-CAL-40217
|  | 13 M㎐ | 3.8 × 10-4
100 pF
|  | 1 k㎐ ~ 4 M㎐ | 3.6 × 10-4
|  | 5 M㎐ | 3.8 × 10-4
|  | 5 M㎐ ~ 10 M㎐ | 4.8 × 10-4
|  | 13 M㎐ | 6.0 × 10-4
1 000 pF
|  | 1 k㎐ ~ 1 M㎐ | 3.6 × 10-4
|  | (1 ~ 4) M㎐ | 5.6 × 10-4
|  | 5 M㎐ | 7.2 × 10-4
|  | 10 M㎐ | 1.9 × 10-3
|  | 13 M㎐ | 2.9 × 10-3
10 ㎋
|  | 120 ㎐ ~ 100 k㎐ | 7.8 × 10-5
100 ㎋
|  | 120 ㎐ ~ 100 k㎐ | 7.8 × 10-5
1 μF
|  | 120 ㎐ | 8.4 × 10-5
|  | 1 k㎐ ~ 10 k㎐ | 7.8 × 10-5
|  | 100 k㎐ | 9.7 × 10-5
10 μF ~ 10 mF
|  | 100 ㎐ | 6.0 × 10-4
|  | 120 ㎐ | 6.0 × 10-4
1 k㎐
|  | 10 μF | 8.0 × 10-4
|  | 100 μF | 8.0 × 10-4
|  | 1 mF | 1.0 × 10-3
|  | 10 mF | 2.4 × 10-3
Resistance |  | 1 mΩ
|  | 1 k㎐ | 5.0 × 10-3
10 mΩ
|  | 1 k㎐ | 5.2 × 10-4
402. Resistance, Capacitance and Inductance
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Impedance bridges/LCR meters | 40217 |  |  | Standard Capacitor Set,
Resistance |  | 100 mΩ |  | Standard Resistor Set,
|  | 1 k㎐ | 3.3 × 10-4 | Standard Inductor/
SS-CAL-40217
1 Ω
|  | 1 k㎐ | 1.2 × 10-4
10 Ω
|  | 1 k㎐ | 1.2 × 10-4
|  | 1 M㎐ | 3.2 × 10-4
|  | 2 M㎐ | 5.2 × 10-4
|  | 3 M㎐ | 6.2 × 10-4
|  | 4 M㎐ | 7.2 × 10-4
|  | 5 M㎐ | 1.0 × 10-3
|  | 10 M㎐ | 4.0 × 10-3
|  | 13 M㎐ | 6.0 × 10-3
100 Ω
|  | 1 k㎐ | 1.3 × 10-4
|  | 1 M㎐ | 3.2 × 10-4
|  | 2 M㎐ | 4.2 × 10-4
|  | 2 M㎐ ~ 5 M㎐ | 5.2 × 10-4
|  | 10 M㎐ | 2.0 × 10-3
|  | 13 M㎐ | 3.0 × 10-3
1 ㏀
|  | 1 k㎐ | 1.2 × 10-4
|  | 100 k㎐ ~ 3 M㎐ | 3.2 × 10-4
|  | 4 M㎐ | 4.2 × 10-4
|  | 5 M㎐ | 5.2 × 10-4
|  | 10 M㎐ | 2.0 × 10-3
|  | 13 M㎐ | 3.0 × 10-3
10 ㏀
|  | 1 k㎐ | 1.2 × 10-4
|  | 100 k㎐ | 2.4 × 10-4
|  | 1 M㎐ | 3.2 × 10-4
100 ㏀
|  | 1 k㎐ | 1.4 × 10-4
|  | 100 k㎐ ~ 1 M㎐ | 3.2 × 10-4
Inductance |  | (1 kHz)
|  | 100 μH | 1.9 × 10-4
|  | 100 μH ~ 10 H | 1.5 × 10-4
DC Bias(±) |  | 0 ~ 100 V | 8.4 × 10-6
DC Current |  | (0 ~ 200) mA | 7.6 × 10-5
|  | 200 mA ~ 2 A | 1.2 × 10-4
|  | (2 ~ 20) A | 4.8 × 10-4
|  | (20 ~ 100) A | 3.0 × 10-4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC ammeters | 40301 |  |  | Calibrator/
AC Current |  | 1 k㎐ |  | SS-CAL-40301
|  | 20 μA | 6.0 × 10-4
(1 ~ 10) k㎐
|  | 20 μA | 5.0 × 10-3
40 ㎐ ~ 1 k㎐
|  | (20 ~ 100) μA | 2.2 × 10-4
|  | 100 μA ~ 1 mA | 1.8 × 10-4
|  | (1 ~ 10) mA | 1.8 × 10-4
|  | (10 ~ 100) mA | 1.8 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 4.0 × 10-4
(1 ~ 5) k㎐
|  | (20 ~ 100) μA | 6.0 × 10-4
|  | 100 μA ~ 1 mA | 4.0 × 10-4
|  | (1 ~ 10) mA | 3.4 × 10-4
|  | (10 ~ 100) mA | 3.2 × 10-4
|  | 100 mA ~ 1 A | 8.0 × 10-4
(5 ~ 10) k㎐
|  | (20 ~ 100) μA | 1.8 × 10-3
|  | 100 μA ~ 1 mA | 1.6 × 10-3
|  | (1 ~ 10) mA | 1.6 × 10-3
|  | (10 ~ 100) mA | 1.0 × 10-3
|  | 100 mA ~ 1 A | 6.4 × 10-3
(50 ~ 60) ㎐
|  | (10 ~ 20) A | 7.0 × 10-4
|  | (20 ~ 30) A | 8.7 × 10-4
|  | (30 ~ 50) A | 7.0 × 10-4
|  | (50 ~ 100) A | 5.3 × 10-4
(60 ~ 100) ㎐
|  | (10 ~ 20) A | 8.0 × 10-4
|  | (20 ~ 30) A | 1.3 × 10-3
|  | (30 ~ 50) A | 1.1 × 10-3
|  | (50 ~ 100) A | 8.0 × 10-4
100 ㎐ ~ 1 k㎐
|  | (10 ~ 20) A | 1.5 × 10-3
|  | (20 ~ 30) A | 4.7 × 10-3
|  | (30 ~ 50) A | 3.5 × 10-3
|  | (50 ~ 100) A | 2.7 × 10-3
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Clamp ammeters/voltmeters | 40302 |  |  | Calibrator, Turn Coil
DC Voltage |  | (±) |  | SS-CAL-40302
|  | 10 mV | 7.4 × 10-5
|  | (10 ~ 100) mV | 1.4 × 10-5
|  | 100 mV ~ 1 V | 6.8 × 10-6
|  | (1 ~ 10) V | 4.6 × 10-6
|  | (10 ~ 100) V | 6.4 × 10-6
|  | (100 ~ 300) V | 1.1 × 10-6
|  | (300 ~ 600) V | 8.7 × 10-6
|  | (600 ~ 1 000) V | 8.2 × 10-6
AC Voltage |  | 40 ㎐
|  | 10 mV | 6.0 × 10-4
|  | (10 ~ 100) mV | 2.0 × 10-4
|  | 100 mV ~ 1 V | 1.3 × 10-4
|  | (1 ~ 10) V | 1.3 × 10-4
|  | (10 ~ 100) V | 1.3 × 10-4
|  | (100 ~ 300) V | 1.3 × 10-4
|  | (300 ~ 600) V | 1.2 × 10-4
|  | (600 ~ 1 000) V | 1.1 × 10-4
40 ㎐ ~ 10 ㎑
|  | 10 mV | 5.6 × 10-4
|  | (10 ~ 100) mV | 1.8 × 10-4
|  | 100 mV ~ 1 V | 6.2 × 10-5
|  | (1 ~ 10) V | 6.4 × 10-5
40 ㎐ ~ 1 ㎑
|  | (10 ~ 100) V | 7.6 × 10-5
|  | (100 ~ 300) V | 1.3 × 10-4
|  | (300 ~ 600) V | 1.1 × 10-4
|  | (600 ~ 1 000) V | 1.1 × 10-4
(1 ~ 10) ㎑
|  | (10 ~ 100) V | 7.8 × 10-5
|  | (100 ~ 300) V | 2.1 × 10-4
|  | (300 ~ 600) V | 2.0 × 10-4
|  | (600 ~ 1 000) V | 2.0 × 10-4
(10 ~ 50) ㎑
|  | 10 mV | 7.0 × 10-4
|  | (10 ~ 100) mV | 3.2 × 10-4
|  | 100 mV ~ 1 V | 1.0 × 10-4
|  | (1 ~ 10) V | 1.0 × 10-4
|  | (10 ~ 100) V | 1.1 × 10-4
|  | (100 ~ 300) V | 7.3 × 10-4
|  | (300 ~ 600) V | 7.3 × 10-4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Clamp ammeters/voltmeters | 40302 |  |  | Calibrator, Turn Coil
AC Voltage |  | (10 ~ 20) ㎑ |  | SS-CAL-40302
|  | (600 ~ 1 000) V | 2.0 × 10-4
(20 ~ 30) ㎑
|  | (600 ~ 1 000) V | 7.0 × 10-4
(50 ~ 100) ㎑
|  | 10 mV | 1.2 × 10-3
|  | (10 ~ 100) mV | 4.2 × 10-4
|  | 100 mV ~ 1 V | 1.6 × 10-4
|  | (1 ~ 10) V | 1.4 × 10-4
|  | (10 ~ 100) V | 2.2 × 10-4
|  | (100 ~ 300) V | 2.8 × 10-3
|  | (300 ~ 600) V | 2.7 × 10-3
Resistance |  | 1 Ω | 4.8 × 10-4
|  | (1 ~ 10) Ω | 4.6 × 10-4
|  | 10 Ω ~ 10 kΩ | 3.2 × 10-4
|  | (10 ~ 100) kΩ | 3.3 × 10-4
|  | 100 kΩ ~ 1 MΩ | 3.8 × 10-4
|  | (1 ~ 10) MΩ | 1.5 × 10-3
|  | (10 ~ 100) MΩ | 5.8 × 10-3
DC Current |  | (±)
|  | 10 μA | 7.4 × 10-4
|  | (10 ~ 100) μA | 1.2 × 10-4
|  | 100 μA ~ 1 mA | 5.4 × 10-5
|  | (1 ~ 10) mA | 5.0 × 10-5
|  | (10 ~ 100) mA | 6.0 × 10-5
|  | 100 mA ~ 1 A | 1.1 × 10-4
|  | (1 ~ 10) A | 1.4 × 10-4
|  | (10 ~ 20) A | 1.4 × 10-4
|  | (20 ~ 50) A | 2.2 × 10-4
|  | (50 ~ 100) A | 1.5 × 10-4
|  | (100 ~ 1 000) A | 3.3 × 10-3
|  | (1 000 ~ 2 500) A | 8.8 × 10-3
AC Current |  | 40 ㎐
|  | 10 μA | 1.4 × 10-3
|  | (10 ~ 100) μA | 3.2 × 10-4
|  | 100 μA ~ 1 mA | 2.4 × 10-4
|  | (1 ~ 10) mA | 2.4 × 10-4
|  | (10 ~ 100) mA | 2.4 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 5.8 × 10-4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Clamp ammeters/voltmeters | 40302 |  |  | Calibrator, Turn Coil
AC Current |  | 40 ㎐ ~ 1 ㎑ |  | SS-CAL-40302
|  | 10 μA | 1.1 × 10-3
|  | (10 ~ 100) μA | 2.6 × 10-4
|  | 100 μA ~ 1 mA | 2.0 × 10-4
|  | (1 ~ 10) mA | 2.0 × 10-4
|  | (10 ~ 100) mA | 2.0 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 5.8 × 10-4
(1 ~ 5) ㎑
|  | 10 μA | 1.7 × 10-3
|  | (10 ~ 100) μA | 5.8 × 10-4
|  | 100 μA ~ 1 mA | 3.8 × 10-4
|  | (1 ~ 10) mA | 3.2 × 10-4
|  | (10 ~ 100) mA | 2.8 × 10-4
|  | 100 mA ~ 1 A | 6.6 × 10-4
|  | (1 ~ 10) A | 1.2 × 10-3
(5 ~ 10) ㎑
|  | 10 μA | 8.8 × 10-3
|  | (10 ~ 100) μA | 2.0 × 10-3
|  | 100 μA ~ 1 mA | 2.0 × 10-3
|  | (1 ~ 10) mA | 7.2 × 10-4
|  | (10 ~ 100) mA | 1.4 × 10-3
|  | 100 mA ~ 1 A | 8.6 × 10-3
|  | (1 ~ 10) A | 4.2 × 10-3
(50 ~ 60) ㎐
|  | (10 ~ 20) A | 7.0 × 10-4
|  | (20 ~ 50) A | 6.0 × 10-4
|  | (50 ~ 100) A | 3.8 × 10-4
|  | (100 ~ 200) A | 4.0 × 10-3
|  | (200 ~ 500) A | 3.6 × 10-3
|  | (500 ~ 1 000) A | 3.8 × 10-3
|  | (1 000 ~ 2 500) A | 8.8 × 10-3
|  | (2 500 ~ 5 000) A | 8.7 × 10-3
60 ㎐ ~ 1 ㎑
|  | (10 ~ 20) A | 1.5 × 10-3
|  | (20 ~ 50) A | 3.1 × 10-3
|  | (50 ~ 100) A | 2.0 × 10-3
(60 ~ 400) ㎐
|  | (100 ~ 200) A | 9.0 × 10-3
|  | (200 ~ 500) A | 8.8 × 10-3
|  | (500 ~ 1 000) A | 9.0 × 10-3
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Clamp ammeters/voltmeters | 40302 |  |  | Calibrator, Turn Coil
AC Current |  | (1 000 ~ 2 500) A | 9.2 × 10-3 | SS-CAL-40302
|  | (2 500 ~ 5 000) A | 8.7 × 10-3
Frequency |  | 10 ㎐ ~ 1 M㎐ | 6.0 × 10-7
Turn Current Coil |  | DC Ratio
|  | 2 | 0.04%
|  | 2 ~ 50 | 0.1%
AC Ratio
(50 ~ 60) ㎐
|  | 2 | 0.11%
|  | 2 ~ 50 | 0.12%
(60 ~ 400) ㎐
|  | 2 | 0.53%
|  | 2 ~ 50 | 0.22%
AC voltage/current calibrators | 40303 |  |  | AC Measurement STD, Calibrator
AC Voltage |  | (20 ~ 40) ㎐ |  | Reference Multimeter, Shunt/
|  | 2 mV | 2.1 × 10-4 | SS-CAL-40303
|  | (2 ~ 20) mV | 6.1 × 10-5
|  | (20 ~ 200) mV | 2.3 × 10-5
|  | 200 mV ~ 2 V | 1.7 × 10-5
40 ㎐ ~ 20 ㎑
|  | 2 mV | 2.1 × 10-4
|  | (2 ~ 20) mV | 4.1 × 10-5
|  | (20 ~ 200) mV | 1.5 × 10-5
|  | 200 mV ~ 2 V | 6.0 × 10-6
(20 ~ 50) ㎑
|  | 2 mV | 2.2 × 10-4
|  | (2 ~ 20) mV | 5.2 × 10-5
|  | (20 ~ 200) mV | 2.4 × 10-5
|  | 200 mV ~ 2 V | 6.2 × 10-6
(50 ~ 100) ㎑
|  | 2 mV | 2.4 × 10-4
|  | (2 ~ 20) mV | 8.7 × 10-5
|  | (20 ~ 200) mV | 4.8 × 10-5
|  | 200 mV ~ 2 V | 2.0 × 10-5
(100 ~ 300) ㎑
|  | 2 mV | 3.3 × 10-4
|  | (2 ~ 20) mV | 2.0 × 10-4
|  | (20 ~ 200) mV | 8.6 × 10-5
|  | 200 mV ~ 2 V | 3.0 × 10-5
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltage/current calibrators | 40303 |  |  | AC Measurement STD, Calibrator
AC Voltage |  | (300 ~ 500) ㎑ |  | Reference Multimeter, Shunt/
|  | 2 mV | 6.4 × 10-4 | SS-CAL-40303
|  | (2 ~ 20) mV | 2.8 × 10-4
|  | (20 ~ 200) mV | 1.3 × 10-4
|  | 200 mV ~ 2 V | 4.7 × 10-5
500 ㎑ ~ 1 ㎒
|  | 2 mV | 1.4 × 10-3
|  | (2 ~ 20) mV | 4.4 × 10-4
|  | (20 ~ 200) mV | 2.5 × 10-4
|  | 200 mV ~ 2 V | 2.5 × 10-4
(20 ~ 40) ㎐
|  | (2 ~ 20) V | 1.8 × 10-5
40 ㎐ ~ 1 ㎑
|  | (2 ~ 20) V | 7.5 × 10-6
(1 ~ 20) ㎑
|  | (2 ~ 20) V | 7.8 × 10-6
(20 ~ 50) ㎑
|  | (2 ~ 20) V | 8.8 × 10-6
(50 ~ 100) ㎑
|  | (2 ~ 20) V | 1.4 × 10-5
(100 ~ 300) ㎑
|  | (2 ~ 20) V | 2.7 × 10-5
(300 ~ 500) ㎑
|  | (2 ~ 20) V | 4.7 × 10-5
500 ㎑ ~ 1 ㎒
|  | (2 ~ 20) V | 2.7 × 10-4
(20 ~ 40) ㎐
|  | (20 ~ 60) V | 1.7 × 10-5
40 ㎐ ~ 1 ㎑
|  | (20 ~ 60) V | 7.7 × 10-6
(1 ~ 20) ㎑
|  | (20 ~ 60) V | 8.0 × 10-6
(20 ~ 50) ㎑
|  | (20 ~ 60) V | 9.7 × 10-6
(50 ~ 100) ㎑
|  | (20 ~ 60) V | 1.4 × 10-5
(100 ~ 300) ㎑
|  | (20 ~ 60) V | 3.7 × 10-5
(20 ~ 40) ㎐
|  | (60 ~ 100) V | 1.7 × 10-5
|  | (100 ~ 200) V | 1.7 × 10-5
|  | (200 ~ 600) V | 1.4 × 10-5
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltage/current calibrators | 40303 |  |  | AC Measurement STD, Calibrator
AC Voltage |  | 40 ㎐ ~ 1 ㎑ |  | Reference Multimeter, Shunt/
|  | (60 ~ 100) V | 9.4 × 10-6 | SS-CAL-40303
|  | (100 ~ 200) V | 9.0 × 10-6
|  | (200 ~ 600) V | 1.5 × 10-5
(1 ~ 20) ㎑
|  | (60 ~ 100) V | 9.8 × 10-6
|  | (100 ~ 200) V | 1.0 × 10-5
|  | (200 ~ 600) V | 1.6 × 10-5
(20 ~ 50) ㎑
|  | (60 ~ 100) V | 1.2 × 10-5
|  | (100 ~ 200) V | 1.2 × 10-5
|  | (200 ~ 600) V | 2.4 × 10-5
(50 ~ 100) ㎑
|  | (60 ~ 100) V | 2.0 × 10-5
|  | (100 ~ 200) V | 2.0 × 10-5
|  | (200 ~ 600) V | 6.7 × 10-5
40 ㎐
|  | (600 ~ 1 000) V | 4.4 × 10-6
40 ㎐ ~ 1 ㎑
|  | (600 ~ 1 000) V | 3.8 × 10-6
(1 ~ 10) ㎑
|  | (600 ~ 1 000) V | 1.6 × 10-5
(10 ~ 20) ㎑
|  | (600 ~ 1 000) V | 2.0 × 10-5
(20 ~ 30) ㎑
|  | (600 ~ 1 000) V | 3.2 × 10-5
AC Current |  | (40 ~ 60) ㎐
|  | 10 μA | 1.8 × 10-4
|  | (10 ~ 100) μA | 1.1 × 10-4
|  | 100 μA ~ 1 mA | 6.2 × 10-5
|  | (1 ~ 10) mA | 1.0 × 10-4
|  | (10 ~ 100) mA | 4.6 × 10-5
|  | 100 mA ~ 1 A | 5.4 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.2 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.8 × 10-5
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltage/current calibrators | 40303 |  |  | AC Measurement STD, Calibrator
AC Current |  | (60 ~ 400) ㎐ |  | Reference Multimeter, Shunt/
|  | 10 μA | 2.2 × 10-4 | SS-CAL-40303
|  | (10 ~ 100) μA | 1.6 × 10-4
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 8.4 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.4 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.4 × 10-5
400 ㎐ ~ 1 k㎐
|  | 10 μA | 1.8 × 10-4
|  | (10 ~ 100) μA | 8.8 × 10-5
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 6.6 × 10-5
|  | (10 ~ 100) mA | 4.4 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
(1 ~ 10）k㎐
|  | 10 μA | 2.1 × 10-4
|  | (10 ~ 100) μA | 1.2 × 10-4
|  | 100 μA ~ 1 mA | 5.8 × 10-5
|  | (1 ~ 10) mA | 7.5 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 6.9 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
Frequency |  | 10 ㎐ | 5.8 × 10-4
|  | 10 ㎐ ~ 1 M㎐ | 5.8 × 10-5
Wattmeter calibrators | 40304 |  |  | AC Measurement STD, Calibrator
AC Voltage |  | 40 ㎐ |  | Reference Multimeter, Shunt,
|  | 2 mV | 2.1 × 10-4 | Power analyzer/
|  | (2 ~ 20) mV | 5.1 × 10-5 | SS-CAL-40304
|  | (20 ~ 200) mV | 1.4 × 10-5
|  | 200 mV ~ 2 V | 7.0 × 10-6
40 ㎐ ~ 20 ㎑
|  | 2 mV | 2.1 × 10-4
|  | (2 ~ 20) mV | 4.1 × 10-5
|  | (20 ~ 200) mV | 1.5 × 10-5
|  | 200 mV ~ 2 V | 6.0 × 10-6
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Wattmeter calibrators | 40304 |  |  | AC Measurement STD, Calibrator
AC Voltage |  | (20 ~ 50) ㎑ |  | Reference Multimeter, Shunt,
|  | 2 mV | 2.2 × 10-4 | Power analyzer/
|  | (2 ~ 20) mV | 5.2 × 10-5 | SS-CAL-40304
|  | (20 ~ 200) mV | 2.4 × 10-5
|  | 200 mV ~ 2 V | 6.2 × 10-6
(50 ~ 100) ㎑
|  | 2 mV | 2.4 × 10-4
|  | (2 ~ 20) mV | 8.7 × 10-5
|  | (20 ~ 200) mV | 4.8 × 10-5
|  | 200 mV ~ 2 V | 2.0 × 10-5
(100 ~ 300) ㎑
|  | 2 mV | 3.3 × 10-4
|  | (2 ~ 20) mV | 2.0 × 10-4
|  | (20 ~ 200) mV | 8.6 × 10-5
|  | 200 mV ~ 2 V | 3.0 × 10-5
(300 ~ 500) ㎑
|  | 2 mV | 6.4 × 10-4
|  | (2 ~ 20) mV | 2.8 × 10-4
|  | (20 ~ 200) mV | 1.3 × 10-4
|  | 200 mV ~ 2 V | 4.7 × 10-5
500 ㎑ ~ 1 ㎒
|  | 2 mV | 1.4 × 10-3
|  | (2 ~ 20) mV | 4.4 × 10-4
|  | (20 ~ 200) mV | 2.5 × 10-4
|  | 200 mV ~ 2 V | 2.5 × 10-4
40 ㎐
|  | (2 ~ 20) V | 7.8 × 10-6
40 ㎐ ~ 1 ㎑
|  | (2 ~ 20) V | 7.5 × 10-6
(1 ~ 20) ㎑
|  | (2 ~ 20) V | 7.8 × 10-6
(20 ~ 50) ㎑
|  | (2 ~ 20) V | 8.8 × 10-6
(50 ~ 100) ㎑
|  | (2 ~ 20) V | 1.4 × 10-5
(100 ~ 300) ㎑
|  | (2 ~ 20) V | 2.7 × 10-5
(300 ~ 500) ㎑
|  | (2 ~ 20) V | 4.7 × 10-5
500 ㎑ ~ 1 ㎒
|  | (2 ~ 20) V | 2.7 × 10-4
40 ㎐
|  | (20 ~ 60) V | 7.7 × 10-6
40 ㎐ ~ 1 ㎑
|  | (20 ~ 60) V | 7.7 × 10-6
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Wattmeter calibrators | 40304 |  |  | AC Measurement STD, Calibrator
AC Voltage |  | (1 ~ 20) ㎑ |  | Reference Multimeter, Shunt,
|  | (20 ~ 60) V | 8.0 × 10-6 | Power analyzer/
|  | (20 ~ 50) ㎑ |  | SS-CAL-40304
|  | (20 ~ 60) V | 9.7 × 10-6
(50 ~ 100) ㎑
|  | (20 ~ 60) V | 1.4 × 10-5
(100 ~ 300) ㎑
|  | (20 ~ 60) V | 3.7 × 10-5
40 ㎐
|  | (60 ~ 100) V | 1.0 × 10-5
|  | (100 ~ 200) V | 1.0 × 10-5
|  | (200 ~ 600) V | 1.4 × 10-5
40 ㎐ ~ 1 ㎑
|  | (60 ~ 100) V | 9.4 × 10-6
|  | (100 ~ 200) V | 9.0 × 10-6
|  | (200 ~ 600) V | 1.5 × 10-5
(1 ~ 20) ㎑
|  | (60 ~ 100) V | 9.8 × 10-6
|  | (100 ~ 200) V | 1.0 × 10-5
|  | (200 ~ 600) V | 1.6 × 10-5
(20 ~ 50) ㎑
|  | (60 ~ 100) V | 1.2 × 10-5
|  | (100 ~ 200) V | 1.2 × 10-5
|  | (200 ~ 600) V | 2.4 × 10-5
(50 ~ 100) ㎑
|  | (60 ~ 100) V | 2.0 × 10-5
|  | (100 ~ 200) V | 2.0 × 10-5
|  | (200 ~ 600) V | 6.7 × 10-5
40 ㎐
|  | (600 ~ 1 000) V | 4.4 × 10-6
40 ㎐ ~ 1 ㎑
|  | (600 ~ 1 000) V | 3.8 × 10-6
(1 ~ 10) ㎑
|  | (600 ~ 1 000) V | 1.6 × 10-5
(10 ~ 20) ㎑
|  | (600 ~ 1 000) V | 2.0 × 10-5
(20 ~ 30) ㎑
|  | (600 ~ 1 000) V | 1.2 × 10-5
AC Current |  | (40 ~ 60) ㎐
|  | 10 μA | 1.8 × 10-4
|  | (10 ~ 100) μA | 1.1 × 10-4
|  | 100 μA ~ 1 mA | 6.2 × 10-5
|  | (1 ~ 10) mA | 1.0 × 10-4
|  | (10 ~ 100) mA | 4.6 × 10-5
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Wattmeter calibrators | 40304 |  |  | AC Measurement STD, Calibrator
AC Current |  | 100 mA ~ 1 A | 5.4 × 10-5 | Reference Multimeter, Shunt,
|  | (1 ~ 10) A | 1.1 × 10-4 | Power analyzer/
|  | (10 ~ 20) A | 7.2 × 10-5 | SS-CAL-40304
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.8 × 10-5
(60 ~ 400) ㎐
|  | 10 μA | 2.2 × 10-4
|  | (10 ~ 100) μA | 1.6 × 10-4
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 8.4 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.4 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.4 × 10-5
400 ㎐ ~ 1 k㎐
|  | 10 μA | 1.8 × 10-4
|  | (10 ~ 100) μA | 8.8 × 10-5
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 6.6 × 10-5
|  | (10 ~ 100) mA | 4.4 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
(1 ~ 10）k㎐
|  | 10 μA | 2.1 × 10-4
|  | (10 ~ 100) μA | 1.2 × 10-4
|  | 100 μA ~ 1 mA | 5.8 × 10-5
|  | (1 ~ 10) mA | 7.5 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 6.9 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
Frequency |  | 10 ㎐ | 5.8 × 10-4
|  | 10 ㎐ ~ 1 M㎐ | 5.8 × 10-5
AC Power |  | (50 ~ 60) ㎐
|  | 3 W | 6.7 × 10-4
|  | (3 ~ 12) W | 5.0 × 10-4
|  | (12 ~ 60) W | 3.3 × 10-4
|  | (60 ~ 120) W | 3.5 × 10-4
|  | (120 ~ 600) W | 3.3 × 10-4
|  | 600 W ~ 2.4 kW | 3.3 × 10-4
|  | (2.4 ~ 4.8) kW | 3.3 × 10-4
|  | (4.8 ~ 12) kW | 3.3 × 10-4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Wattmeter calibrators | 40304 |  |  | AC Measurement STD, Calibrator
Power Factor |  | (50 ~ 60) ㎐ |  | Reference Multimeter, Shunt,
|  | -1 ~ 1 | 0.000 22 | Power analyzer/
SS-CAL-40304
Harmonic Voltage |  | (50 ~ 3 000) ㎐
|  | (0.5 ~ 20) % | 0.088%
Harmonic Current |  | (50 ~ 3 000) ㎐
|  | (0.5 ~ 20) % | 0.078%
AC current shunts | 40305 |  |  | Reference Multimeter,
AC Resistance |  | (50 ㎐ ~ 60 ㎐) |  | Calibrator/
|  | 1 mΩ | 4.2 × 10-4 | SS-CAL-40305
(60 ㎐ ~ 100 ㎐)
|  | 1 mΩ | 6.2 × 10-4
(100 ㎐ ~ 1 k㎐)
|  | 1 mΩ | 2.0 × 10-3
(40 ㎐ ~ 1 k㎐）
|  | (1 ~ 10) mΩ | 6.0 × 10-4
|  | (10 ~ 100) mΩ | 4.0 × 10-4
|  | 100 mΩ ~ 1 Ω | 2.6 × 10-4
|  | (1 ~ 10) Ω | 2.6 × 10-4
|  | (10 ~ 100) Ω | 2.6 × 10-4
|  | 100 Ω ~ 1 kΩ | 3.0 × 10-4
Voltage current phase meter | 40307 |  |  | Power Calibrator/
Phase |  | (50 ~ 60) ㎐ |  | SS-CAL-40307
|  | ±(0˚~ 90˚) | 0.1˚
Power factor meters | 40310 |  |  | Power Calibrator/
Power Factor |  | (50 ~ 60) ㎐ |  | SS-CAL-40310
|  | -1 ~ 1 | 0.000 12
AC power meters | 40311 |  |  | Power Calibrator,
DC Voltage |  | (±) |  | Calibrator/
|  | 10 mV | 7.4 × 10-5 | SS-CAL-40311
|  | (10 ~ 100) mV | 1.4 × 10-5
|  | 100 mV ~ 1 V | 6.8 × 10-6
|  | (1 ~ 10) V | 4.6 × 10-6
|  | (10 ~ 100) V | 6.4 × 10-6
|  | (100 ~ 300) V | 1.1 × 10-5
|  | (300 ~ 600) V | 8.7 × 10-6
|  | (600 ~ 1 000) V | 8.2 × 10-6
AC Voltage |  | 40 ㎐
|  | 10 mV | 6.0 × 10-4
|  | (10 ~ 100) mV | 2.0 × 10-4
|  | 100 mV ~ 1 V | 1.3 × 10-4
|  | (1 ~ 10) V | 1.3 × 10-4
|  | (10 ~ 100) V | 1.3 × 10-4
|  | (100 ~ 300) V | 1.3 × 10-4
|  | (300 ~ 600) V | 1.2 × 10-4
|  | (600 ~ 1 000) V | 1.1 × 10-4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311 |  |  | Power Calibrator,
AC Voltage |  | 40 ㎐ ~ 10 ㎑ |  | Calibrator/
|  | 10 mV | 5.6 × 10-4 | SS-CAL-40311
|  | (10 ~ 100) mV | 1.8 × 10-4
|  | 100 mV ~ 1 V | 6.2 × 10-5
|  | (1 ~ 10) V | 6.4 × 10-5
40 ㎐ ~ 1 ㎑
|  | (10 ~ 100) V | 7.6 × 10-5
|  | (100 ~ 300) V | 1.3 × 10-4
|  | (300 ~ 600) V | 1.1 × 10-4
|  | (600 ~ 1 000) V | 1.1 × 10-4
(1 ~ 10) ㎑
|  | (10 ~ 100) V | 7.8 × 10-5
|  | (100 ~ 300) V | 2.1 × 10-4
|  | (300 ~ 600) V | 2.0 × 10-4
|  | (600 ~ 1 000) V | 2.0 × 10-4
(10 ~ 50) ㎑
|  | 10 mV | 7.0 × 10-4
|  | (10 ~ 100) mV | 3.2 × 10-4
|  | 100 mV ~ 1 V | 1.0 × 10-4
|  | (1 ~ 10) V | 1.0 × 10-4
|  | (10 ~ 100) V | 1.1 × 10-4
|  | (100 ~ 300) V | 7.3 × 10-4
|  | (300 ~ 600) V | 7.3 × 10-4
(10 ~ 20) ㎑
|  | (600 ~ 1 000) V | 2.0 × 10-4
(20 ~ 30) ㎑
|  | (600 ~ 1 000) V | 7.0 × 10-4
(50 ~ 100) ㎑
|  | 10 mV | 1.2 × 10-3
|  | (10 ~ 100) mV | 4.2 × 10-4
|  | 100 mV ~ 1 V | 1.6 × 10-4
|  | (1 ~ 10) V | 1.4 × 10-4
|  | (10 ~ 100) V | 2.2 × 10-4
|  | (100 ~ 300) V | 2.8 × 10-3
|  | (300 ~ 600) V | 2.7 × 10-3
DC Current |  | (±)
|  | 10 μA | 7.4 × 10-4
|  | (10 ~ 100) μA | 1.2 × 10-4
|  | 100 μA ~ 1 mA | 5.4 × 10-5
|  | (1 ~ 10) mA | 5.0 × 10-5
|  | (10 ~ 100) mA | 6.0 × 10-5
|  | 100 mA ~ 1 A | 1.1 × 10-4
|  | (1 ~ 10) A | 1.4 × 10-4
|  | (10 ~ 20) A | 1.4 × 10-4
|  | (20 ~ 50) A | 2.2 × 10-4
|  | (50 ~ 100) A | 1.5 × 10-4
DC(Clamp Type) |  | (100 ~ 1 000) A | 3.3 × 10-3
|  | (1 000 ~ 2 500) A | 8.8 × 10-3
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311 |  |  | Power Calibrator,
AC Current |  | 40 ㎐ |  | Calibrator/
|  | 10 μA | 1.4 × 10-3 | SS-CAL-40311
|  | (10 ~ 100) μA | 3.2 × 10-4
|  | 100 μA ~ 1 mA | 2.4 × 10-4
|  | (1 ~ 10) mA | 2.4 × 10-4
|  | (10 ~ 100) mA | 2.4 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 5.8 × 10-4
40 ㎐ ~ 1 ㎑
|  | 10 μA | 1.1 × 10-3
|  | (10 ~ 100) μA | 2.6 × 10-4
|  | 100 μA ~ 1 mA | 2.0 × 10-4
|  | (1 ~ 10) mA | 2.0 × 10-4
|  | (10 ~ 100) mA | 2.0 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 5.8 × 10-4
(1 ~ 5) ㎑
|  | 10 μA | 1.7 × 10-3
|  | (10 ~ 100) μA | 5.8 × 10-4
|  | 100 μA ~ 1 mA | 3.8 × 10-4
|  | (1 ~ 10) mA | 3.2 × 10-4
|  | (10 ~ 100) mA | 2.8 × 10-4
|  | 100 mA ~ 1 A | 6.6 × 10-4
|  | (1 ~ 10) A | 1.2 × 10-3
(5 ~ 10) ㎑
|  | 10 μA | 8.8 × 10-3
|  | (10 ~ 100) μA | 2.0 × 10-3
|  | 100 μA ~ 1 mA | 2.0 × 10-3
|  | (1 ~ 10) mA | 7.2 × 10-4
|  | (10 ~ 100) mA | 1.4 × 10-3
|  | 100 mA ~ 1 A | 8.6 × 10-3
|  | (1 ~ 10) A | 4.2 × 10-3
(50 ~ 60) ㎐
|  | (10 ~ 20) A | 7.0 × 10-4
|  | (20 ~ 50) A | 6.0 × 10-4
|  | (50 ~ 100) A | 3.8 × 10-4
AC(Clamp Type) |  | (100 ~ 200) A | 4.0 × 10-3
|  | (200 ~ 500) A | 3.6 × 10-3
|  | (500 ~ 1 000) A | 3.8 × 10-3
|  | (1 000 ~ 2 500) A | 8.8 × 10-3
|  | (2 500 ~ 5 000) A | 8.7 × 10-3
60 ㎐ ~ 1 ㎑
|  | (10 ~ 20) A | 1.5 × 10-3
|  | (20 ~ 50) A | 3.1 × 10-3
|  | (50 ~ 100) A | 2.0 × 10-3
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311 |  |  | Power Calibrator,
AC Current |  | (60 ~ 400) ㎐ |  | Calibrator/
AC(Clamp Type) |  | (100 ~ 200) A | 9.0 × 10-3 | SS-CAL-40311
|  | (200 ~ 500) A | 8.8 × 10-3
|  | (500 ~ 1 000) A | 9.0 × 10-3
|  | (1 000 ~ 2 500) A | 9.2 × 10-3
|  | (2 500 ~ 5 000) A | 8.7 × 10-3
Frequency |  | 10 ㎐ ~ 1 M㎐ | 6.0 × 10-7
DC Power |  | 1 mW ~ 48 W | 1.8 × 10-4
|  | (48 ~ 120) W | 3.2 × 10-4
|  | (120 ~ 240) W | 2.8 × 10-4
|  | (240 ~ 480) W | 4.6 × 10-4
|  | (480 ~ 600) W | 6.9 × 10-4
|  | 600 W~ 1.2 kW | 7.0 × 10-4
|  | (1.2 ~ 2.4) kW | 6.3 × 10-4
|  | (2.4 ~ 5) kW | 6.4 × 10-4
|  | (5 ~ 7) kW | 6.3 × 10-4
|  | (7 ~ 20) kW | 1.2 × 10-3
AC Power |  | (50 ~ 60) ㎐
|  | 6 mW | 1.0 × 10-2
|  | 6 mW ~ 6 W | 1.7 × 10-4
|  | 6 W ~ 2.4 kW | 1.2 × 10-4
|  | 2.4 kW ~ 4.8 kW | 1.3 × 10-4
|  | (4.8 ~ 24) kW | 1.2 × 10-4
|  | (24 ~ 35.2) kW | 1.9 × 10-4
Power Factor |  | (50 ~ 60) ㎐
|  | -1 ~ 1 | 0.000 12
Harmonic Voltage |  | (50 ~ 3 000) ㎐
|  | (0.5 ~ 20) % | 0.06%
Harmonic Current |  | (50 ~ 3 000) ㎐
|  | (0.5 ~ 20) % | 0.052%
AC power supplies | 40312 |  |  | Voltage Standard, Multimeter,
AC Voltage |  | 40 ㎐ |  | Current Shunt/
|  | 100 mV | 2.0 × 10-4 | SS-CAL-40312
|  | 100 mV ~ 100 V | 1.6 × 10-4
|  | (100 ~ 300) V | 2.1 × 10-4
|  | (300 ~ 500) V | 1.8 × 10-4
|  | (500 ~ 1 000) V | 1.6 × 10-4
40 ㎐ ~ 1 k㎐
|  | 100 mV | 1.6 × 10-4
|  | 100 mV ~ 100 V | 1.0 × 10-4
|  | (100 ~ 300) V | 1.9 × 10-4
|  | (300 ~ 1 000) V | 1.6 × 10-4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power supplies | 40312 |  |  | Voltage Standard, Multimeter,
AC Current |  | 40 ㎐ |  | Current Shunt/
|  | 200 mA | 3.3 × 10-4 | SS-CAL-40312
|  | 200 mA ~ 2 A | 4.5 × 10-4
|  | (2 ~ 20) A | 6.2 × 10-4
|  | (20 ~ 50) A | 6.5 × 10-4
40 ㎐ ~ 1 k㎐
|  | 200 mA | 3.0 × 10-4
|  | 200 mA ~ 2 A | 4.3 × 10-4
|  | (2 ~ 20) A | 6.0 × 10-4
|  | (20 ~ 50) A | 2.2 × 10-3
Frequency |  | 10 ㎐ | 5.8 × 10-4
|  | 10 ㎐ ~ 10 k㎐ | 1.2 × 10-4
Puncture/safety testers | 40313 |  |  | AC/DC Kilovoltmeter,
DC Voltage |  | (±) |  | High Voltage Digital Meter,
|  | 0.1 kV | 6.0 × 10-3 | Reference Multimeter,Counter/
|  | (0.1 ~ 100) kV | 3.0 × 10-3 | SS-CAL-40313
AC Voltage |  | (50 ~ 60) ㎐
|  | 0.1 kV | 1.2 × 10-2
|  | (0.1 ~ 100) kV | 1.0 × 10-2
DC Current |  | 0.1 mA | 5.8 × 10-3
|  | (0.1 ~ 0.5) mA | 1.2 × 10-3
|  | (0.5 ~ 1) mA | 5.8 × 10-4
|  | (1 ~ 2) mA | 2.9 × 10-3
|  | (2 ~ 5) mA | 1.2 × 10-3
|  | (5 ~ 10) mA | 5.8 × 10-4
|  | (10 ~ 20) mA | 3.0 × 10-4
|  | (20 ~ 50) mA | 1.2 × 10-4
|  | (50 ~ 100) mA | 5.8 × 10-4
AC Current |  | (50 ~ 60) ㎐
|  | 0.1 mA | 5.8 × 10-3
|  | (0.1 ~ 0.5) mA | 1.4 × 10-3
|  | (0.5 ~ 1) mA | 7.8 × 10-4
|  | (1 ~ 2) mA | 3.2 × 10-3
|  | (2 ~ 5) mA | 1.4 × 10-3
|  | (5 ~ 10) mA | 7.8 × 10-4
|  | (10 ~ 20) mA | 9.1 × 10-4
|  | (20 ~ 50) mA | 5.2 × 10-4
|  | (50 ~ 100) mA | 7.0 × 10-4
Timer |  | 0.1 s | 5.8 × 10-3
|  | (0.1 ~ 0.2) s | 2.9 × 10-3
|  | (0.2 ~ 0.5) s | 1.2 × 10-3
|  | (0.5 ~ 1) s | 5.8 × 10-4
|  | (1 ~ 2) s | 2.9 × 10-4
|  | (2 ~ 5) s | 1.2 × 10-4
|  | (5 ~ 10) s | 5.8 × 10-5
|  | (10 ~ 20) s | 2.9 × 10-5
|  | (20 ~ 50) s | 1.2 × 10-5
|  | (50 ~ 100) s | 5.8 × 10-6
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltmeters | 40318 |  |  | Reference Multimeter,
AC Voltage |  | 40 ㎐ ~ 20 k㎐ |  | Calibrator/
|  | 2 mV | 2.5 × 10-3 | SS-CAL-40318
|  | (2 ~ 20) mV | 3.1 × 10-4
|  | (20 ~ 200) mV | 1.0 × 10-4
|  | 200 mV ~ 2 V | 3.8 × 10-5
|  | (2 ~ 20) V | 3.8 × 10-5
|  | (20 ~ 200) V | 5.0 × 10-5
(20 ~ 50) k㎐
|  | 2 mV | 2.4 × 10-3
|  | (2 ~ 20) mV | 3.2 × 10-4
|  | (20 ~ 200) mV | 1.3 × 10-4
|  | 200 mV ~ 2 V | 7.0 × 10-5
|  | (2 ~ 20) V | 7.0 × 10-5
|  | (20 ~ 200) V | 8.0 × 10-5
(50 ~ 100) k㎐
|  | 2 mV | 3.1 × 10-3
|  | (2 ~ 20) mV | 5.0 × 10-4
|  | (20 ~ 200) mV | 3.0 × 10-4
|  | 200 mV ~ 2 V | 1.1 × 10-4
|  | (2 ~ 20) V | 1.1 × 10-4
|  | (20 ~ 200) V | 1.5 × 10-4
(100 ~ 300) k㎐
|  | 2 mV | 6.1 × 10-3
|  | (2 ~ 20) mV | 9.0 × 10-4
|  | (20 ~ 200) mV | 4.6 × 10-4
|  | 200 mV ~ 2 V | 3.3 × 10-4
|  | (2 ~ 20) V | 3.4 × 10-4
(300 ~ 500) k㎐
|  | 2 mV | 1.2 × 10-2
|  | (2 ~ 20) mV | 1.9 × 10-3
|  | (20 ~ 200) mV | 9.0 × 10-4
|  | 200 mV ~ 2 V | 9.0 × 10-4
|  | (2 ~ 20) V | 9.0 × 10-4
500 k㎐ ~ 1 M㎐
|  | 2 mV | 1.4 × 10-2
|  | (2 ~ 20) mV | 3.8 × 10-3
|  | (20 ~ 200) mV | 2.8 × 10-3
|  | 200 mV ~ 2 V | 1.4 × 10-3
|  | (2 ~ 20) V | 1.7 × 10-3
(1 ~ 30) M㎐
|  | (2 ~ 20) mV | 3.0 × 10-3
|  | (20 ~ 200) mV | 3.0 × 10-3
|  | 200 mV ~ 2 V | 3.0 × 10-3
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltmeters | 40318 |  |  | Reference Multimeter,
AC Voltage |  | 40 ㎐ ~ 1 k㎐ |  | Calibrator/
|  | (200 ~ 600) V | 6.7 × 10-5 | SS-CAL-40318
|  | (600 ~ 1 000) V | 6.8 × 10-5
(1 ~ 20) k㎐
|  | (200 ~ 600) V | 1.1 × 10-4
|  | (600 ~ 1 000) V | 1.0 × 10-4
(20 ~ 50) k㎐
|  | (200 ~ 600) V | 2.3 × 10-4
(50 ~ 100) k㎐
|  | (200 ~ 600) V | 5.3 × 10-4
(20 ~ 30) k㎐
|  | (600 ~ 1 000) V | 2.0 × 10-4
DC Output Voltage |  | (±)
|  | 10 mV | 6.0 × 10-5
|  | (1 ~ 100) mV | 8.0 × 10-6
|  | 100 mV ~ 1 V | 6.2 × 10-5
AC Output Voltage |  | 40 ㎐
|  | 10 mV | 6.0 × 10-4
|  | (10 ~ 100) mV | 1.6 × 10-4
|  | 100 mV ~ 1 V | 1.4 × 10-4
40 ㎐ ~ 1 k㎐
|  | 10 mV | 4.0 × 10-4
|  | (10 ~ 100) mV | 1.4 × 10-4
|  | 100 mV ~ 1 V | 1.1 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF amplifiers | 40401 |  |  | Reference Multimeter/
Amplifier |  | (40 ㎐) |  | SS-CAL-40401
|  | (0 ~ 10) ㏈ | 0.01 ㏈
|  | (10 ~ 20) ㏈ | 0.02 ㏈
|  | (20 ~ 30) ㏈ | 0.06 ㏈
|  | (30 ~ 40) ㏈ | 0.16 ㏈
|  | (40 ~ 50) ㏈ | 0.61 ㏈
|  | (50 ~ 60) ㏈ | 1.6 ㏈
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF amplifiers | 40401 |  |  | Reference Multimeter/
Amplifier |  | (40 ~ 100) ㎐ |  | SS-CAL-40401
|  | (0 ~ 10) ㏈ | 0.01 ㏈
|  | (10 ~ 20) ㏈ | 0.01 ㏈
|  | (20 ~ 30) ㏈ | 0.05 ㏈
|  | (30 ~ 40) ㏈ | 0.10 ㏈
|  | (40 ~ 50) ㏈ | 0.45 ㏈
|  | (50 ~ 60) ㏈ | 1.2 ㏈
100 ㎐ ~ 1 k㎐
|  | (0 ~ 10) ㏈ | 0.01 ㏈
|  | (10 ~ 20) ㏈ | 0.01 ㏈
|  | (20 ~ 30) ㏈ | 0.04 ㏈
|  | (30 ~ 40) ㏈ | 0.09 ㏈
|  | (40 ~ 50) ㏈ | 0.49 ㏈
|  | (50 ~ 60) ㏈ | 1.1 ㏈
(1 ~ 10) k㎐
|  | (0 ~ 10) ㏈ | 0.01 ㏈
|  | (10 ~ 20) ㏈ | 0.01 ㏈
|  | (20 ~ 30) ㏈ | 0.05 ㏈
|  | (30 ~ 40) ㏈ | 0.11 ㏈
(10 ~ 100) k㎐
|  | (0 ~ 10) ㏈ | 0.04 ㏈
|  | (10 ~ 20) ㏈ | 0.07 ㏈
|  | (20 ~ 30) ㏈ | 0.36 ㏈
|  | (30 ~ 40) ㏈ | 0.68 ㏈
DC/LF attenuators | 40402 | (－) |  | Reference Multimeter/
Attenuation |  | (40 ㎐) |  | SS-CAL-40402
|  | (0 ~ 20) ㏈ | 1.0 × 10-3
|  | (20 ~ 30) ㏈ | 3.2 × 10-4
|  | (30 ~ 50) ㏈ | 3.2 × 10-3
|  | (50 ~ 60) ㏈ | 1.0 × 10-2
(40 ~ 100) ㎐
|  | (0 ~ 20) ㏈ | 1.0 × 10-3
|  | (20 ~ 30) ㏈ | 3.2 × 10-4
|  | (30 ~ 50) ㏈ | 3.2 × 10-3
|  | (50 ~ 60) ㏈ | 1.0 × 10-2
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC/LF attenuators | 40402 |  |  | Reference Multimeter/
Attenuation |  | (100 ㎐ ~ 1 k㎐) |  | SS-CAL-40402
|  | (0 ~ 20) ㏈ | 1.0 × 10-3
|  | (20 ~ 30) ㏈ | 3.2 × 10-4
|  | (30 ~ 50) ㏈ | 3.2 × 10-3
|  | (50 ~ 60) ㏈ | 1.0 × 10-2
(1 ~ 10) k㎐
|  | (0 ~ 20) ㏈ | 1.0 × 10-3
|  | (20 ~ 30) ㏈ | 3.2 × 10-4
|  | (30 ~ 50) ㏈ | 3.2 × 10-3
|  | (50 ~ 60) ㏈ | 1.0 × 10-2
(10 ~ 100) k㎐
|  | (0 ~ 20) ㏈ | 1.3 × 10-3
|  | (20 ~ 30) ㏈ | 1.6 × 10-3
|  | (30 ~ 50) ㏈ | 9.5 × 10-3
|  | (50 ~ 60) ㏈ | 2.0 × 10-2
Multimeter calibrators | 40403 |  |  | Standard Resistor, Calibrator
DC Voltage |  | (±) |  | Reference Multimeter, Shunt,
|  | 1 mV | 3.8 × 10-4 | AC Measurement STD/
|  | (1 ~ 10) mV | 3.8 × 10-5 | SS-CAL-40403
|  | (10 ~ 100) mV | 3.6 × 10-6
|  | 100 mV ~ 1 V | 7.0 × 10-7
|  | (1 ~ 10) V | 5.0 × 10-7
|  | (10 ~ 100) V | 6.0 × 10-7
|  | (100 ~ 1 000) V | 8.0 × 10-7
AC Voltage |  | 40 ㎐
|  | 2 mV | 2.1 × 10-4
|  | (2 ~ 20) mV | 5.1 × 10-5
|  | (20 ~ 200) mV | 1.4 × 10-5
|  | 200 mV ~ 2 V | 7.0 × 10-6
40 ㎐ ~ 20 ㎑
|  | 2 mV | 2.1 × 10-4
|  | (2 ~ 20) mV | 4.1 × 10-5
|  | (20 ~ 200) mV | 1.5 × 10-5
|  | 200 mV ~ 2 V | 6.0 × 10-6
(20 ~ 50) ㎑
|  | 2 mV | 2.2 × 10-4
|  | (2 ~ 20) mV | 5.2 × 10-5
|  | (20 ~ 200) mV | 2.4 × 10-5
|  | 200 mV ~ 2 V | 6.2 × 10-6
(50 ~ 100) ㎑
|  | 2 mV | 2.4 × 10-4
|  | (2 ~ 20) mV | 8.7 × 10-5
|  | (20 ~ 200) mV | 4.8 × 10-5
|  | 200 mV ~ 2 V | 2.0 × 10-5
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Multimeter calibrators | 40403 |  |  | Standard Resistor, Calibrator
AC Voltage |  | (100 ~ 300) ㎑ |  | Reference Multimeter, Shunt,
|  | 2 mV | 3.3 × 10-4 | AC Measurement STD/
|  | (2 ~ 20) mV | 2.0 × 10-4 | SS-CAL-40403
|  | (20 ~ 200) mV | 8.6 × 10-5
|  | 200 mV ~ 2 V | 3.0 × 10-5
(300 ~ 500) ㎑
|  | 2 mV | 6.4 × 10-4
|  | (2 ~ 20) mV | 2.8 × 10-4
|  | (20 ~ 200) mV | 1.3 × 10-4
|  | 200 mV ~ 2 V | 4.7 × 10-5
500 ㎑ ~ 1 ㎒
|  | 2 mV | 1.4 × 10-3
|  | (2 ~ 20) mV | 4.4 × 10-4
|  | (20 ~ 200) mV | 2.5 × 10-4
|  | 200 mV ~ 2 V | 2.5 × 10-4
40 ㎐
|  | (2 ~ 20) V | 7.8 × 10-6
40 ㎐ ~ 1 ㎑
|  | (2 ~ 20) V | 7.5 × 10-6
(1 ~ 20) ㎑
|  | (2 ~ 20) V | 7.8 × 10-6
(20 ~ 50) ㎑
|  | (2 ~ 20) V | 8.8 × 10-6
(50 ~ 100) ㎑
|  | (2 ~ 20) V | 1.4 × 10-5
(100 ~ 300) ㎑
|  | (2 ~ 20) V | 2.7 × 10-5
(300 ~ 500) ㎑
|  | (2 ~ 20) V | 4.7 × 10-5
500 ㎑ ~ 1 ㎒
|  | (2 ~ 20) V | 2.7 × 10-4
40 ㎐
|  | (20 ~ 60) V | 7.7 × 10-6
40 ㎐ ~ 1 ㎑
|  | (20 ~ 60) V | 7.7 × 10-6
(1 ~ 20) ㎑
|  | (20 ~ 60) V | 8.0 × 10-6
(20 ~ 50) ㎑
|  | (20 ~ 60) V | 9.7 × 10-6
(50 ~ 100) ㎑
|  | (20 ~ 60) V | 1.4 × 10-5
(100 ~ 300) ㎑
|  | (20 ~ 60) V | 3.7 × 10-5
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Multimeter calibrators | 40403 |  |  | Standard Resistor, Calibrator
AC Voltage |  | 40 ㎐ |  | Reference Multimeter, Shunt,
|  | (60 ~ 100) V | 1.0 × 10-5 | AC Measurement STD/
|  | (100 ~ 200) V | 1.0 × 10-5 | SS-CAL-40403
|  | (200 ~ 600) V | 1.4 × 10-5
40 ㎐ ~ 1 ㎑
|  | (60 ~ 100) V | 9.4 × 10-6
|  | (100 ~ 200) V | 9.0 × 10-6
|  | (200 ~ 600) V | 1.5 × 10-5
(1 ~ 20) ㎑
|  | (60 ~ 100) V | 9.8 × 10-6
|  | (100 ~ 200) V | 1.0 × 10-5
|  | (200 ~ 600) V | 1.6 × 10-5
(20 ~ 50) ㎑
|  | (60 ~ 100) V | 1.2 × 10-5
|  | (100 ~ 200) V | 1.2 × 10-5
|  | (200 ~ 600) V | 2.4 × 10-5
(50 ~ 100) ㎑
|  | (60 ~ 100) V | 2.0 × 10-5
|  | (100 ~ 200) V | 2.0 × 10-5
|  | (200 ~ 600) V | 6.7 × 10-5
40 ㎐
|  | (600 ~ 1 000) V | 1.5 × 10-5
40 ㎐ ~ 1 ㎑
|  | (600 ~ 1 000) V | 1.5 × 10-5
(1 ~ 10) ㎑
|  | (600 ~ 1 000) V | 1.6 × 10-5
(10 ~ 20) ㎑
|  | (600 ~ 1 000) V | 2.0 × 10-5
(20 ~ 30) ㎑
|  | (600 ~ 1 000) V | 3.2 × 10-5
DC Current |  | (±)
|  | 10 μA | 7.0 × 10-6
|  | 10 μA ~ 100 mA | 3.8 × 10-6
|  | 100 mA ~ 1 A | 4.2 × 10-6
|  | (1 ~ 10) A | 6.2 × 10-6
|  | (10 ~ 20) A | 2.5 × 10-5
|  | (20 ~ 50) A | 1.2 × 10-5
|  | (50 ~ 100) A | 3.0 × 10-5
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Multimeter calibrators | 40403 |  |  | Standard Resistor, Calibrator
AC Current |  | (40 ~ 60) ㎐ |  | Reference Multimeter, Shunt,
|  | 10 μA | 1.8 × 10-4 | AC Measurement STD/
|  | (10 ~ 100) μA | 1.1 × 10-4 | SS-CAL-40403
|  | 100 μA ~ 1 mA | 6.2 × 10-5
|  | (1 ~ 10) mA | 1.0 × 10-4
|  | (10 ~ 100) mA | 4.6 × 10-5
|  | 100 mA ~ 1 A | 5.4 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.2 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.8 × 10-5
(60 ~ 400) ㎐
|  | 10 μA | 2.2 × 10-4
|  | (10 ~ 100) μA | 1.6 × 10-4
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 8.4 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.4 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.4 × 10-5
400 ㎐ ~ 1 k㎐
|  | 10 μA | 1.8 × 10-4
|  | (10 ~ 100) μA | 8.8 × 10-5
|  | 100 μA ~ 1 mA | 4.8 × 10-5
|  | (1 ~ 10) mA | 6.6 × 10-5
|  | (10 ~ 100) mA | 4.4 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
(1 ~ 10）k㎐
|  | 10 μA | 2.1 × 10-4
|  | (10 ~ 100) μA | 1.2 × 10-4
|  | 100 μA ~ 1 mA | 5.8 × 10-5
|  | (1 ~ 10) mA | 7.5 × 10-5
|  | (10 ~ 100) mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 5.2 × 10-5
|  | (1 ~ 10) A | 1.1 × 10-4
|  | (10 ~ 20) A | 6.9 × 10-5
|  | (20 ~ 50) A | 8.8 × 10-5
|  | (50 ~ 100) A | 7.6 × 10-5
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Multimeter calibrators | 40403 |  |  | Standard Resistor, Calibrator
Resistance |  | 1 Ω | 3.2 × 10-5 | Reference Multimeter, Shunt,
|  | (1 ~ 1.9) Ω | 2.3 × 10-5 | AC Measurement STD/
|  | (1.9 ~ 10) Ω | 5.4 × 10-5 | SS-CAL-40403
|  | (10 ~ 19) Ω | 4.5 × 10-6
|  | (19 ~ 100) Ω | 2.0 × 10-6
|  | (100 ~ 190) Ω | 2.1 × 10-6
|  | 190 Ω ~ 1 kΩ | 2.0 × 10-6
|  | (1 ~ 1.9) kΩ | 2.4 × 10-6
|  | (1.9 ~ 10) kΩ | 1.9 × 10-6
|  | (10 ~ 19) kΩ | 2.4 × 10-6
|  | (19 ~ 100) kΩ | 2.0 × 10-6
|  | (100 ~ 190) kΩ | 2.9 × 10-6
|  | 190 kΩ ~ 1 MΩ | 2.6 × 10-6
|  | (1 ~ 1.9) MΩ | 4.8 × 10-6
|  | (1.9 ~ 10) MΩ | 1.1 × 10-5
|  | (10 ~ 19) MΩ | 2.0 × 10-5
|  | (19 ~ 100) MΩ | 4.8 × 10-5
Frequency |  | 10 ㎐ | 5.8 × 10-4
|  | 10 ㎐ ~ 1 M㎐ | 5.8 × 10-5
Oscilloscope calibrators | 40404 |  |  | Calibrator/
Impedance Mesurement |  | 50 Ω | 2.8 × 10-4 | SS-CAL-40404
|  | 1 MΩ | 2.0 × 10-4
DC Voltage Amplitude |  | (±)
|  | 1 mV | 6.0 × 10-4
|  | (1 ~ 10) mV | 8.0 × 10-5
|  | 10 mV ~ 100 V | 6.0 × 10-5
|  | (100 ~ 200) V | 2.9 × 10-5
AC Voltage Amplitude |  | 100 ㎐~ 10 k㎐
|  | 1 mV | 1.4 × 10-3
|  | (1 ~ 10) mV | 1.4 × 10-4
|  | 10 mV ~ 100 V | 6.0 × 10-5
|  | (100 ~ 200) V | 4.0 × 10-5
Sine Wave Generator |  | 50 k㎐ | 2.3 × 10-2
|  | 50 k㎐ ~ 100 M㎐ | 2.3 × 10-2
|  | 100 M㎐ ~ 2 G㎐ | 2.3 × 10-2
|  | 2 ~ 6 G㎐ | 3.0 × 10-2
Time Marker Generator |  | 0.1 ns | 6.0 × 10-6
|  | 0.1 ns ~ 0.2 ns | 3.0 × 10-6
|  | 0.2 ns ~ 0.5 ns | 1.2 × 10-6
|  | 0.5 ns ~ 5 s | 6.0 × 10-7
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Audio distortion analyzers/meters | 40407 |  |  | Calibrator/
AC Voltage |  | (1 mV) |  | SS-CAL-40407
|  | 40 ㎐ ~ 10 k㎐ | 5.0 × 10-3
|  | (10 ~ 100) k㎐ | 7.0 × 10-3
(1 ~ 10) mV
|  | 40 ㎐ ~ 10 k㎐ | 8.0 × 10-4
|  | (10 ~ 100) k㎐ | 1.0 × 10-3
(10 ~ 100) mV
|  | 40 ㎐ ~ 10 k㎐ | 6.4 × 10-4
|  | (10 ~ 100) k㎐ | 7.4 × 10-4
(100 mV ~100 V)
|  | 40 ㎐ ~ 100 k㎐ | 6.0 × 10-4
(100 ~ 1 000) V
|  | 100 ㎐ ~ 1 k㎐ | 6.0 × 10-4
DC Voltage |  | (1 ~ 10) mV | 1.0 × 10-3
|  | 10 mV ~ 100 V | 6.0 × 10-4
|  | 100 V ~ 500 V | 2.0 × 10-4
D/T Calibrator |  | (400 ㎐ , 1 k㎐)
Distortion |  | (-10 ~ -50) ㏈ | 0.26 ㏈
|  | (-50 ~ -60) ㏈ | 0.30 ㏈
|  | (-60 ~ -70) ㏈ | 0.62 ㏈
|  | (-70 ~ -80) ㏈ | 1.2 ㏈
LF filters | 40408 |  |  | Audio Analyzer/
Filter |  | 10.0 ㎐ ~ 100.0 ㎑ | 1.0 × 10-3 | SS-CAL-40408
LF/Audiosignalanalyzers | 40409 |  |  | Calibrator, Reference
Output Frequency |  | 1 ㎐ ~ 200 k㎐ | 5.8 × 10-6 | Multimeter/
SS-CAL-40409
AC Output Level |  | 40 ㎐
|  | 10 mV | 2.8 × 10-4
|  | (10 ~ 100) mV | 2.3 × 10-4
|  | 100 mV ~ 1 V | 2.3 × 10-4
|  | (1 ~ 10) V | 2.9 × 10-4
|  | (10 ~ 30) V | 7.0 × 10-4
|  | (10 ~ -20) ㏈m | 6.0 × 10-3 ~㏈m
(40 ~ 100) ㎐
|  | 10 mV | 2.4 × 10-4
|  | (10 ~ 100) mV | 1.9 × 10-4
|  | 100 mV ~ 1 V | 1.9 × 10-4
|  | (1 ~ 10) V | 2.6 × 10-4
|  | (10 ~ 30) V | 6.6 × 10-4
|  | (10 ~ -20) ㏈m | 6.0 × 10-3 ~㏈m
(100 ㎐ ~ 1 k㎐)
|  | 10 mV | 2.0 × 10-4
|  | (10 ~ 100) mV | 1.6 × 10-4
|  | 100 mV ~ 1 V | 1.9 × 10-4
|  | (1 ~ 10) V | 2.3 × 10-4
|  | (10 ~ 30) V | 4.4 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF/Audiosignalanalyzers | 40409 |  |  | Calibrator, Reference
AC Output Level |  | (10 ~ -20) ㏈m | 6.0 × 10-3 ~㏈m | Multimeter/
SS-CAL-40409
(1 ~ 10) k㎐
|  | 10 mV | 2.4 × 10-4
|  | (10 ~ 100) mV | 1.9 × 10-4
|  | 100 mV ~ 1 V | 1.9 × 10-4
|  | (1 ~ 10) V | 2.6 × 10-4
|  | (10 ~ 30) V | 6.6 × 10-4
|  | (10 ~ -20) ㏈m | 6.0 × 10-3 ~㏈m
(10 ~ 100) k㎐
|  | 10 mV | 1.8 × 10-3
|  | (10 ~ 100) mV | 1.2 × 10-3
|  | 100 mV ~ 1 V | 1.2 × 10-3
|  | (1 ~ 10) V | 1.6 × 10-3
|  | (10 ~ 30) V | 3.8 × 10-3
|  | (10 ~ -20) ㏈m | 1.8 × 10-3 ~㏈m
AC Output Level Flatness |  | 1 k㎐, 0 ㏈ (Ref)
|  | 40 ㎐ | 1.4 × 10-3
|  | (40 ~ 100) ㎐ | 1.2 × 10-3
|  | 100 ㎐ ~ 10 k㎐ | 1.2 × 10-3
|  | (10 ~ 100) k㎐ | 7.2 × 10-3
Output Attenuation |  | 1 k㎐, 0 ㏈ (Ref)
|  | (0 ~ -60) ㏈ | 5.8 × 10-3 ~㏈
Output DC Offset |  | 0 mV ~ 10 mV | 0.6 μV
|  | 10 mV ~ 10 V | 6.4 × 10-5
|  | 10 V ~ 50 V | 1.5 × 10-5
Input Frequency |  | 1 ㎐ ~ 200 k㎐ | 6.2 × 10-7
AC Input Level Flatness |  | 1 k㎐, 0 ㏈ (Ref)
|  | 40 ㎐ | 2.7 × 10-3
|  | 40 ㎐ ~ 10 k㎐ | 7.9 × 10-4
|  | (10 ~ 100) k㎐ | 1.4 × 10-3
DC Input Level |  | 0 mV | 0.6 μV
|  | 0 mV ~ 100 V | 5.8 × 10-4
|  | (100 ~ 300) V | 1.9 × 10-4
Input Distortion |  | (400 ㎐, 1 k㎐)
|  | -20 ㏈ | 0.25 ㏈
|  | (-20 ~ -50) ㏈ | 0.25 ㏈
|  | (-50 ~ -60) ㏈ | 0.30 ㏈
|  | (-60 ~ -70) ㏈ | 1.1 ㏈
|  | (-70 ~ -80) ㏈ | 1.1 ㏈
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF/Audiosignalanalyzers | 40409 |  |  | Calibrator, Reference
AC Input Level |  | (40 ㎐) |  | Multimeter/
|  | 1 mV | 1.7 × 10-3 | SS-CAL-40409
|  | (1 ~ 10) mV | 3.4 × 10-4
|  | (10 ~ 100) mV | 3.0 × 10-4
|  | 100 mV ~ 1 V | 7.4 × 10-4
|  | (1 ~ 10) V | 3.4 × 10-4
|  | (10 ~ 100) V | 3.4 × 10-4
|  | (100 ~ 300) V | 4.5 × 10-4
(40 ㎐ ~ 100 ㎐)
|  | 1 mV | 8.4 × 10-4
|  | (1 ~ 10) mV | 1.4 × 10-4
|  | (10 ~ 100) mV | 9.4 × 10-5
|  | 100 mV ~ 1 V | 7.4 × 10-5
|  | (1 ~ 10) V | 7.2 × 10-5
|  | (10 ~ 100) V | 8.4 × 10-5
|  | (100 ~ 300) V | 2.1 × 10-4
(100 ㎐ ~ 1 k㎐)
|  | 1 mV | 8.4 × 10-4
|  | (1 ~ 10) mV | 1.4 × 10-4
|  | (10 ~ 100) mV | 9.4 × 10-5
|  | 100 mV ~ 1 V | 7.4 × 10-5
|  | (1 ~ 10) V | 7.2 × 10-5
|  | (10 ~ 100) V | 8.4 × 10-5
|  | (100 ~ 300) V | 2.1 × 10-4
(1 k㎐ ~ 10 k㎐)
|  | 1 mV | 7.8 × 10-4
|  | (1 ~ 10) mV | 1.4 × 10-4
|  | (10 ~ 100) mV | 9.2 × 10-5
|  | 100 mV ~ 1 V | 7.4 × 10-5
|  | (1 ~ 10) V | 7.2 × 10-5
|  | (10 ~ 100) V | 8.4 × 10-5
(10 k㎐ ~ 100 k㎐)
|  | 1 mV | 1.0 × 10-3
|  | (1 ~ 10) mV | 6.2 × 10-4
|  | (10 ~ 100) mV | 5.6 × 10-4
|  | 100 mV ~ 1 V | 1.8 × 10-4
|  | (1 ~ 10) V | 1.5 × 10-4
|  | (10 ~ 100) V | 2.4 × 10-4
Input Attenuation |  | (1 k㎐)
|  | (30 ~ -30) ㏈ | 5.9 × 10-3 ~ ㏈
|  | (-30 ~ -50) ㏈ | 6.9 × 10-3 ~ ㏈
|  | (-50 ~ -60) ㏈ | 2.0 × 10-2 ~ ㏈
Line frequency meters | 40410 |  |  | Calibrator/
Frequency |  | 10 ㎐ | 1.0 × 10-3 | SS-CAL-40410
|  | 10 ㎐ ~ 10 k㎐ | 2.0 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Function generators | 40411 |  |  | Audio Analyzer, Digital
Frequency |  | 1 ㎐ ~ 1 G㎐ | 5.8 × 10-9 | Multimeter/
|  | (1 ~ 4) G㎐ | 1.5 × 10-8 | SS-CAL-40411
Output Level |  | (40 ㎐)
|  | 1 mV | 4.8 × 10-3
|  | (1 ~ 10) mV | 6.0 × 10-4
|  | (1 ~ 100) mV | 1.9 × 10-4
|  | 10 mV ~ 10 V | 1.5 × 10-4
|  | (10 ~ 100) V | 1.4 × 10-4
(40 ㎐ ~ 100 ㎐)
|  | 1 mV | 4.8 × 10-3
|  | (1 ~ 10) mV | 5.8 × 10-3
|  | (10 ~ 100) mV | 1.7 × 10-4
|  | 100 mV ~ 100 V | 1.2 × 10-4
100 ㎐ ~ 1 k㎐
|  | 1 mV | 2.5 × 10-3
|  | (1 ~ 10) mV | 3.6 × 10-4
|  | (10 ~ 100) mV | 1.4 × 10-4
|  | 100 mV ~ 100 V | 9.8 × 10-5
(1 k㎐ ~ 10 k㎐)
|  | 1 mV | 4.8 × 10-3
|  | (1 ~ 10) mV | 5.8 × 10-4
|  | (10 ~ 100) mV | 1.7 × 10-4
|  | 100 mV ~ 100 V | 1.2 × 10-4
(10 k㎐ ~ 100 k㎐)
|  | 1 mV | 2.4 × 10-2
|  | (1 ~ 10) mV | 3.2 × 10-3
|  | (10 ~ 100) mV | 1.0 × 10-3
|  | 100 mV ~ 100 V | 8.2 × 10-4
DC Offset |  | (-20 ~ 20) V | 6.0 × 10-5
Sine Wave Flatness |  | (40 ㎐ ~ 10 k㎐)
|  | 1 k㎐,0 ㏈ Ref | 0.006 ㏈
100 mV ~ 30 V
(10 ~ 100) k㎐
1 k㎐,0 ㏈ Ref
|  | 100 mV | 0.011 ㏈
|  | 100 mV ~ 1 V | 0.009 ㏈
|  | (1 ~ 30) V | 0.013 ㏈
Attenuation |  | (100 ㎐ ~ 100 k㎐)
|  | (10 ~ -60)㏈ | 0.13 ㏈
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Function generators | 40411 |  |  | Audio Analyzer, Digital
Distortion |  | 20 ㎐ ~ 100 k㎐ |  | Multimeter/
|  | (10.0 ~ 0.010) % | 6.6 × 10-2 | SS-CAL-40411
Rise Fall Time |  | 10 μs ~ 1 ns | 1.2 × 10-3
|  | 1 ns ~ 100 ㎰ | 1.2 × 10-3
Duty cycle |  | (1 ~ 99) % | 0.006 2 %
AC/DC high voltages volt meters | 40413 |  |  | Calibrator/
DC Voltage |  | (±) |  | SS-CAL-40413
|  | 0.1 kV | 1.0 × 10-2
|  | (0.1 ~ 100) kV | 3.0 × 10-3
AC Voltage |  | 60 ㎐
|  | 0.1 kV | 2.0 × 10-2
|  | (0.1 ~ 100) kV | 1.1 × 10-2
Leakage current testers | 40416 |  |  | Calibrator/
AC Voltage |  | 40 ㎐ ~ 1 k㎐ |  | SS-CAL-40416
|  | 100 mV | 2.0 × 10-4
|  | 100 mV ~ 1 000 V | 1.0 × 10-4
AC Current |  | 1 k㎐
|  | 20 μA | 1.0 × 10-3
40 ㎐ ~ 1 k㎐
|  | 20 μA ~ 10 mA | 6.0 × 10-4
|  | (10 ~ 100) mA | 2.0 × 10-4
(±)
DC Current |  | 10 μA | 1.2 × 10-3
|  | 10 μA ~ 10 mA | 6.0 × 10-4
|  | (10 ~ 100) mA | 1.0 × 10-4
Resistance |  | 1 Ω ~ 10 kΩ | 6.0 × 10-5
Electronic AC/DC loads | 40417 |  |  | Calibrator/
DC Voltage |  | 100 mV | 6.0 × 10-5 | SS-CAL-40417
|  | 100 mV ~ 1 V | 6.0 × 10-5
|  | (1 ~ 100) V | 8.0 × 10-6
|  | (100 ~ 500) V | 1.6 × 10-5
|  | (500 ~ 1 000) V | 6.0 × 10-5
DC Current |  | 100 ㎃ | 1.0 × 10-4
|  | 100 ㎃ ~ 1 A | 1.4 × 10-4
|  | (1 ~ 10) A | 6.0 × 10-4
|  | (10 ~ 20) A | 5.0 × 10-4
|  | (20 ~ 100) A | 3.2 × 10-4
|  | (100 ~ 200) A | 5.0 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electronic AC/DC loads | 40417 |  |  | Calibrator/
AC Voltage |  | (50 ~ 60) ㎐ |  | SS-CAL-40417
|  | 1 V | 1.0 × 10-4
|  | (1 ~ 10) V | 6.4 × 10-5
|  | (10 ~ 100) V | 7.8 × 10-5
|  | (100 ~ 300) V | 1.1 × 10-4
|  | (300 ~ 600) V | 9.3 × 10-5
AC Current |  | (50 ~ 60) ㎐
|  | 100 mA | 4.2 × 10-4
|  | 100 mA ~ 1 A | 5.2 × 10-4
` |  | (1 ~ 5) A | 8.0 × 10-4
|  | (5 ~ 10) A | 6.8 × 10-4
|  | (10 ~ 20) A | 7.0 × 10-4
|  | (20 ~ 40) A | 8.5 × 10-4
Charge & Discharge tester
Sense Voltage (Meter) |  | 100 mV | 6.0 × 10-5
|  | 100 mV ~ 1 V | 8.0 × 10-5
|  | (1 ~ 10) V | 8.0 × 10-6
|  | (10 ~ 100) V | 1.0 × 10-5
|  | (100 ~ 500) V | 8.0 × 10-6
|  | (500 ~ 1 000) V | 6.0 × 10-5
DC Voltage (Source) |  | 100 mV | 6.0 × 10-5
|  | 100 mV ~ 1 V | 6.0 × 10-5
|  | (1 ~ 10) V | 8.0 × 10-6
|  | (10 ~ 100) V | 1.0 × 10-5
|  | (100 ~ 500) V | 8.0 × 10-6
|  | (500 ~ 1 000) V | 8.0 × 10-5
DC Current (Charge) |  | 100 ㎃ | 1.4 × 10-4
|  | 100 ㎃ ~ 1 A | 2.0 × 10-4
|  | (1 ~ 10) A | 6.0 × 10-4
|  | (10 ~ 20) A | 5.0 × 10-4
|  | (20 ~ 100) A | 3.2 × 10-4
|  | (100 ~ 200) A | 3.8 × 10-4
|  | (200 ~ 300) A | 4.0 × 10-4
|  | (300 ~ 600) A | 4.0 × 10-4
|  | (600 ~ 1 000) A | 3.8 × 10-4
DC Current (Discharge) |  | -100 ㎃ | 1.2 × 10-4
|  | -100 ㎃ ~ -1 A | 2.0 × 10-4
|  | (-1 ~ -10) A | 6.0 × 10-4
|  | (-10 ~ -20) A | 5.0 × 10-4
|  | (-20 ~ -100) A | 3.2 × 10-4
|  | (-100 ~ -200) A | 3.8 × 10-4
|  | (-200 ~ -300) A | 4.0 × 10-4
|  | (-300 ~ -600) A | 4.0 × 10-4
|  | (-600 ~ -1 000) A | 3.8 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Analogue/Digital multimeters | 40419 |  |  | Calibrator/
DC Voltage |  | (±) |  | SS-CAL-40419
|  | 1 mV | 6.8 × 10-4
|  | (1 ~ 10) mV | 7.0 × 10-5
|  | (10 ~ 100) mV | 8.2 × 10-6
|  | 100 mV ~ 1 V | 3.4 × 10-6
|  | (1 ~ 10) V | 2.0 × 10-6
|  | (10 ~ 100) V | 2.8 × 10-6
|  | (100 ~ 1 000) V | 3.2 × 10-6
AC Voltage |  | 10 ㎐
|  | 2 mV | 2.6 × 10-3
|  | (2 ~ 20) mV | 4.7 × 10-4
|  | (20 ~ 200) mV | 3.0 × 10-4
|  | 200 mV ~ 2 V | 2.6 × 10-4
|  | (2 ~ 20) V | 2.6 × 10-4
|  | (20 ~ 200) V | 2.6 × 10-4
40 ㎐ ~ 20 k㎐
|  | 2 mV | 2.5 × 10-3
|  | (2 ~ 20) mV | 3.1 × 10-4
|  | (20 ~ 200) mV | 1.0 × 10-4
|  | 200 mV ~ 2 V | 3.8 × 10-5
|  | (2 ~ 20) V | 3.8 × 10-5
|  | (20 ~ 200) V | 5.0 × 10-5
(20 ~ 50) k㎐
|  | 2 mV | 2.4 × 10-3
|  | (2 ~ 20) mV | 3.2 × 10-4
|  | (20 ~ 200) mV | 1.3 × 10-4
|  | 200 mV ~ 2 V | 7.0 × 10-5
|  | (2 ~ 20) V | 7.0 × 10-5
|  | (20 ~ 200) V | 8.0 × 10-5
(50 ~ 100) k㎐
|  | 2 mV | 3.1 × 10-3
|  | (2 ~ 20) mV | 5.0 × 10-4
|  | (20 ~ 200) mV | 3.0 × 10-4
|  | 200 mV ~ 2 V | 1.1 × 10-4
|  | (2 ~ 20) V | 1.1 × 10-4
|  | (20 ~ 200) V | 1.5 × 10-4
(100 ~ 300) k㎐
|  | 2 mV | 6.1 × 10-3
|  | (2 ~ 20) mV | 9.0 × 10-3
|  | (20 ~ 200) mV | 4.6 × 10-4
|  | 200 mV ~ 2 V | 3.3 × 10-4
|  | (2 ~ 20) V | 3.4 × 10-4
(300 ~ 500) k㎐
|  | 2 mV | 1.2 × 10-2
|  | (2 ~ 20) mV | 1.9 × 10-3
|  | (20 ~ 200) mV | 9.0 × 10-4
|  | 200 mV ~ 2 V | 9.0 × 10-4
|  | (2 ~ 20) V | 9.0 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Analogue/Digital multimeters | 40419 |  |  | Calibrator/
AC Voltage |  | 500 k㎐ ~ 1 M㎐ |  | SS-CAL-40419
|  | 2 mV | 1.4 × 10-2
|  | (2 ~ 20) mV | 3.8 × 10-3
|  | (20 ~ 200) mV | 2.8 × 10-3
|  | 200 mV ~ 2 V | 1.4 × 10-3
|  | (2 ~ 20) V | 1.7 × 10-3
40 ㎐ ~ 1 k㎐
|  | (200 ~ 600) V | 7.3 × 10-5
|  | (600 ~ 1 000) V | 6.8 × 10-5
(1 ~ 20) k㎐
|  | (200 ~ 600) V | 1.1 × 10-4
|  | (600 ~ 1 000) V | 1.0 × 10-4
(20 ~ 50) k㎐
|  | (200 ~ 600) V | 2.3 × 10-4
(50 ~ 100) k㎐
|  | (200 ~ 600) V | 5.3 × 10-4
(20 ~ 30) k㎐
|  | (600 ~ 1 000) V | 2.0 × 10-4
Resistance |  | 1 Ω | 8.0 × 10-7
|  | (1 ~ 10) Ω | 8.0 × 10-7
|  | 10 Ω ~ 100 Ω | 6.0 × 10-7
|  | 100 Ω ~ 1 kΩ | 6.0 × 10-7
|  | (1 ~ 10) kΩ | 4.0 × 10-7
|  | (10 ~ 100) kΩ | 6.0 × 10-7
|  | 100 kΩ ~ 1 MΩ | 1.2 × 10-6
|  | (1 ~ 10) MΩ | 6.0 × 10-6
|  | (10 ~ 100) MΩ | 9.2 × 10-6
|  | 100 MΩ ~ 1 GΩ | 6.0 × 10-5
DC Current |  | (±)
|  | 1 nA | 1.4 × 10-2
|  | (1 ~ 100) nA | 8.4 × 10-3
|  | 100 nA ~ 1 μA | 2.4 × 10-3
|  | (1 ~ 10) μA | 2.0 × 10-4
|  | (10 ~ 100) μA | 4.8 × 10-5
|  | 100 μA ~ 100 mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 6.6 × 10-5
|  | (1 ~ 20) A | 1.9 × 10-4
|  | (20 ~ 30) A | 3.3 × 10-4
AC Current |  | 1 k㎐
|  | 20 μA | 6.0 × 10-4
(1 ~ 10) k㎐
|  | 20 μA | 5.0 × 10-3
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Analogue/Digital multimeters | 40419 |  |  | Calibrator/
AC Current |  | 10 ㎐ |  | SS-CAL-40419
|  | (20 ~ 100) μA | 4.6 × 10-4
|  | 100 μA ~ 1 mA | 3.4 × 10-4
|  | (1 ~ 10) mA | 4.0 × 10-4
|  | (10 ~ 100) mA | 4.0 × 10-4
40 ㎐ ~ 1 k㎐
|  | (20 ~ 100) μA | 2.2 × 10-4
|  | 100 μA ~ 1 mA | 1.8 × 10-4
|  | (1 ~ 10) mA | 1.8 × 10-4
|  | (10 ~ 100) mA | 1.8 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 4.0 × 10-4
(1 ~ 5) k㎐
|  | (20 ~ 100) μA | 6.0 × 10-4
|  | 100 μA ~ 1 mA | 4.0 × 10-4
|  | (1 ~ 10) mA | 3.4 × 10-4
|  | (10 ~ 100) mA | 3.2 × 10-4
|  | 100 mA ~ 1 A | 8.0 × 10-4
(5 ~ 10) k㎐
|  | (20 ~ 100) μA | 1.8 × 10-3
|  | 100 μA ~ 1 mA | 1.6 × 10-3
|  | (1 ~ 10) mA | 1.6 × 10-3
|  | (10 ~ 100) mA | 1.0 × 10-3
|  | 100 mA ~ 1 A | 6.4 × 10-3
(50 ~ 60) ㎐
|  | (10 ~ 20) A | 7.0 × 10-4
|  | (20 ~ 30) A | 8.7 × 10-4
(60 ~ 100) ㎐
|  | (10 ~ 20) A | 8.0 × 10-4
|  | (20 ~ 30) A | 1.3 × 10-3
100 ㎐ ~ 1 k㎐
|  | (10 ~ 20) A | 1.5 × 10-3
|  | (20 ~ 30) A | 4.7 × 10-3
Frequency |  | 10 ㎐ ~ 10 M㎐ | 1.0 × 10-6
Noise meters | 40420 |  |  | Calibrator/
AC Voltage Test |  | (40 ㎐ ~ 100 k㎐) |  | SS-CAL-40420
|  | 1 mV | 6.4 × 10-3
|  | (1 ~ 3) mV | 1.9 × 10-3
|  | (3 ~ 10) mV | 7.2 × 10-4
|  | (10 ~ 30) mV | 1.3 × 10-3
|  | (30 ~ 100) mV | 7.4 × 10-4
|  | (100 ~ 300) mV | 2.7 × 10-4
|  | 300 mV ~ 1 V | 2.0 × 10-4
|  | (1 ~ 3) V | 2.3 × 10-4
|  | (3 ~ 10) V | 1.6 × 10-4
|  | (10 ~ 30) V | 4.3 × 10-4
|  | (30 ~ 100) V | 2.4 × 10-4
(40 ㎐ ~ 1 k㎐)
|  | (100 ~ 1 000) V | 1.0 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Noise meters | 40420 |  |  | Calibrator/
AC Voltage Test |  | (1 k㎐ ~ 1 M㎐) |  | SS-CAL-40420
|  | (1 ~ 10) mV | 6.2 × 10-3
|  | (10 ~ 100) mV | 4.0 × 10-3
|  | 100 mV ~ 1 V | 2.4 × 10-3
|  | (1 ~ 20) V | 2.8 × 10-3
(1 ㎒ ~ 30 M㎐)
|  | 25 mV | 8.0 × 10-3
|  | (25 ~ 100) mV | 7.6 × 10-3
|  | (100 ~ 300) mV | 7.3 × 10-3
|  | 300 mV ~ 1 V | 5.4 × 10-3
|  | (1 ~ 2) V | 5.0 × 10-3
AC Voltage Output |  | (10 ㎐ ~ 100 k㎐)
|  | (1 ~ 100) mV | 1.4 × 10-4
|  | 100 mV ~ 1 V | 1.0 × 10-4
DC Voltage Output |  | (±)
|  | (1 ~ 100) mV | 8.0 × 10-6
|  | 100 mV ~ 1 V | 1.0 × 10-5
Oscilloscopes | 40421 |  |  | Calibration Generator/
Impedance Measure |  | 50 Ω | 8.0 × 10-6 | SS-CAL-40421
|  | (50 ~ 75) Ω | 8.0 × 10-6
|  | 75 Ω ~ 1 MΩ | 2.0 × 10-5
DC Voltage |  | (±)
|  | 1 mV | 8.0 × 10-4
|  | (1 ~ 10) mV | 1.0 × 10-4
|  | (10 ~ 100) mV | 1.0 × 10-5
|  | 100 mV ~ 1 V | 6.2 × 10-5
|  | (1 ~ 10) V | 6.0 × 10-6
|  | (10 ~ 100) V | 8.0 × 10-6
|  | (100 ~ 120) V | 6.7 × 10-5
|  | (120 ~ 200) V | 4.0 × 10-5
AC Voltage |  | at 1 k㎐
|  | 2 mV | 2.4 × 10-2
|  | (2 ~ 3) mV | 1.7 × 10-2
|  | (3 ~ 5) mV | 1.0 × 10-2
|  | (5 ~ 6) mV | 9.0 × 10-3
|  | (6 ~ 10) mV | 5.8 × 10-3
|  | (10 ~ 12) mV | 5.0 × 10-3
|  | (12 ~ 20) mV | 3.5 × 10-3
|  | (20 ~ 30) mV | 2.7 × 10-3
|  | (30 ~ 60) mV | 2.0 × 10-3
|  | (60 ~ 100) mV | 1.6 × 10-3
|  | 100 mV ~ 0.12 V | 1.7 × 10-3
|  | (0.12 ~ 0.2) V | 1.5 × 10-3
|  | (0.2 ~ 120) V | 1.3 × 10-3
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Oscilloscopes | 40421 |  |  | Calibration Generator/
Sine Wave Signal Generator Level |  | 50 k㎐ | 2.7 × 10-2 | SS-CAL-40421
|  | 50 k㎐ ~ 18 G㎐ | 3.3 × 10-2
|  | (18 ~ 26.5) G㎐ | 4.0 × 10-2
Time Marker |  | 2 ns | 3.0 × 10-3
|  | 2 ns ~ 20 ms | 6.0 × 10-4
|  | (20 ~ 50) ms | 1.5 × 10-4
|  | (50 ~ 100) ms | 6.2 × 10-4
|  | (100 ~ 200) ms | 4.0 × 10-4
|  | (200 ~ 500) ms | 6.4 × 10-4
|  | 500 ms ~ 1 s | 1.4 × 10-3
|  | (1 ~ 2) s | 2.4 × 10-3
|  | (2 ~ 5) s | 6.0 × 10-3
CAL Output Amplitude |  | DC Voltage
|  | 100 mV | 4.0 × 10-6
|  | 100 mV ~ 10 V | 8.0 × 10-6
AC Voltage
(100 ㎐ ~ 10 k㎐)
|  | 100 mV | 2.6 × 10-4
|  | 100 mV ~ 10 V | 1.8× 10-4
CAL Output Frequency |  | 100 ㎐ | 6.0 × 10-7
|  | 100 ㎐ ~ 10 M㎐ | 6.0 × 10-7
Volt/Current recorders | 40424 | (±) |  | Calibrator/
DC Voltage |  | 1 mV | 6.8 × 10-4 | SS-CAL-40424
|  | (1 ~ 10) mV | 7.0 × 10-5
|  | (10 ~ 100) mV | 8.2 × 10-6
|  | 100 mV ~ 1 V | 3.4 × 10-6
|  | (1 ~ 10) V | 2.0 × 10-6
|  | (10 ~ 100) V | 2.8 × 10-6
|  | (100 ~ 1 000) V | 3.2 × 10-6
AC Voltage |  | 40 ㎐ ~ 20 k㎐
|  | 2 mV | 2.5 × 10-3
|  | (2 ~ 20) mV | 3.1 × 10-4
|  | (20 ~ 200) mV | 1.0 × 10-4
|  | 200 mV ~ 2 V | 3.8 × 10-5
|  | (2 ~ 20) V | 3.8 × 10-5
|  | (20 ~ 200) V | 5.0 × 10-5
(20 ~ 50) k㎐
|  | 2 mV | 2.4 × 10-3
|  | (2 ~ 20) mV | 3.2 × 10-4
|  | (20 ~ 200) mV | 1.3 × 10-4
|  | 200 mV ~ 2 V | 7.0 × 10-5
|  | (2 ~ 20) V | 7.0 × 10-5
|  | (20 ~ 200) V | 8.0 × 10-5
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Volt/Current recorders | 40424 |  |  | Calibrator/
AC Voltage |  | (50 ~ 100) k㎐ |  | SS-CAL-40424
|  | 2 mV | 3.1 × 10-3
|  | (2 ~ 20) mV | 5.0 × 10-4
|  | (20 ~ 200) mV | 3.0 × 10-4
|  | 200 mV ~ 2 V | 1.1 × 10-4
|  | (2 ~ 20) V | 1.1 × 10-4
|  | (20 ~ 200) V | 1.5 × 10-4
(100 ~ 300) k㎐
|  | 2 mV | 6.1 × 10-3
|  | (2 ~ 20) mV | 9.0 × 10-4
|  | (20 ~ 200) mV | 4.6 × 10-4
|  | 200 mV ~ 2 V | 3.3 × 10-4
|  | (2 ~ 20) V | 3.4 × 10-4
(300 ~ 500) k㎐
|  | 2 mV | 1.2 × 10-2
|  | (2 ~ 20) mV | 1.9 × 10-3
|  | (20 ~ 200) mV | 9.0 × 10-4
|  | 200 mV ~ 2 V | 9.0 × 10-4
|  | (2 ~ 20) V | 9.0 × 10-4
500 k㎐ ~ 1 M㎐
|  | 2 mV | 1.4 × 10-2
|  | (2 ~ 20) mV | 3.8 × 10-3
|  | (20 ~ 200) mV | 2.8 × 10-3
|  | 200 mV ~ 2 V | 1.4 × 10-3
|  | (2 ~ 20) V | 1.7 × 10-3
40 ㎐ ~ 1 k㎐
|  | (200 ~ 600) V | 7.3 × 10-5
|  | (600 ~ 1 000) V | 6.8 × 10-5
(1 ~ 20) k㎐
|  | (200 ~ 600) V | 1.1 × 10-4
|  | (600 ~ 1 000) V | 1.0 × 10-4
(20 ~ 50) k㎐
|  | (200 ~ 600) V | 2.3 × 10-4
(50 ~ 100) k㎐
|  | (200 ~ 600) V | 5.3 × 10-4
(20 ~ 30) k㎐
|  | (600 ~ 1 000) V | 2.0 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Volt/Current recorders | 40424 |  |  | Calibrator/
DC Current |  | (±) |  | SS-CAL-40424
|  | 1 nA | 1.4 × 10-2
|  | (1 ~ 100) nA | 8.4 × 10-3
|  | 100 nA ~ 1 μA | 2.4 × 10-3
|  | (1 ~ 10) μA | 2.0 × 10-4
|  | (10 ~ 100) μA | 4.8 × 10-5
|  | 100 μA ~ 100 mA | 4.2 × 10-5
|  | 100 mA ~ 1 A | 6.6 × 10-5
|  | (1 ~ 20) A | 1.9 × 10-4
|  | (20 ~ 30) A | 3.3 × 10-4
|  | (30 ~ 50) A | 2.4 × 10-4
|  | (50 ~ 100) A | 1.6 × 10-4
AC Current |  | 1 k㎐
|  | 20 μA | 6.0 × 10-4
(1 ~ 10) k㎐
|  | 20 μA | 5.0 × 10-3
40 ㎐ ~ 1 k㎐
|  | (20 ~ 100) μA | 2.2 × 10-4
|  | 100 μA ~ 1 mA | 1.8 × 10-4
|  | (1 ~ 10) mA | 1.8 × 10-4
|  | (10 ~ 100) mA | 1.8 × 10-4
|  | 100 mA ~ 1 A | 3.6 × 10-4
|  | (1 ~ 10) A | 4.0 × 10-4
(1 ~ 5) k㎐
|  | (20 ~ 100) μA | 6.0 × 10-4
|  | 100 μA ~ 1 mA | 4.0 × 10-4
|  | (1 ~ 10) mA | 3.4 × 10-4
|  | (10 ~ 100) mA | 3.2 × 10-4
|  | 100 mA ~ 1 A | 8.0 × 10-4
(5 ~ 10) k㎐
|  | (20 ~ 100) μA | 1.8 × 10-3
|  | 100 μA ~ 1 mA | 1.6 × 10-3
|  | (1 ~ 10) mA | 1.6 × 10-3
|  | (10 ~ 100) mA | 1.0 × 10-3
|  | 100 mA ~ 1 A | 6.4 × 10-3
(50 ~ 60) ㎐
|  | (10 ~ 20) A | 7.0 × 10-4
|  | (20 ~ 30) A | 8.7 × 10-4
|  | (30 ~ 50) A | 6.0 × 10-4
|  | (50 ~ 100) A | 3.8 × 10-4
(60 ~ 100) ㎐
|  | (10 ~ 20) A | 8.0 × 10-4
|  | (20 ~ 30) A | 1.3 × 10-3
|  | (30 ~ 50) A | 9.2 × 10-4
|  | (50 ~ 100) A | 6.0 × 10-4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Volt/Current recorders | 40424 |  |  | Calibrator/
AC Current |  | 100 ㎐ ~ 1 k㎐ |  | SS-CAL-40424
|  | (10 ~ 20) A | 1.5 × 10-3
|  | (20 ~ 30) A | 4.7 × 10-3
|  | (30 ~ 50) A | 3.2 × 10-3
|  | (50 ~ 100) A | 2.0 × 10-3
Relay test sets | 40425 | (±) |  | Digital Multimeter/
DC Voltage |  | 10 mV | 6.0 × 10-5 | SS-CAL-40425
|  | (10 ~ 100) mV | 8.0 × 10-6
|  | 100 mV ~ 10 V | 6.8 × 10-6
|  | (10 ~ 100) V | 8.0 × 10-6
|  | (100 ~ 300) V | 6.7 × 10-6
|  | (300 ~ 500) V | 1.2 × 10-5
|  | (500 ~ 1 000) V | 8.0 × 10-6
DC Current |  | 1 mA | 6.0 × 10-4
|  | (1 ~ 10) mA | 8.0 × 10-5
|  | (10 ~ 200) mA | 7.0 × 10-5
|  | 200 mA ~ 2 A | 1.2 × 10-4
|  | (2 ~ 20) A | 4.8 × 10-4
|  | (20 ~ 100) A | 2.0 × 10-4
|  | (100 ~ 200) A | 4.0 × 10-4
AC Voltage |  | 40 ㎐
|  | 100 mV | 2.0 × 10-4
|  | 100 mV ~ 100 V | 1.6 × 10-4
|  | (100 ~ 300) V | 2.1 × 10-4
|  | (300 ~ 500) V | 1.8 × 10-4
|  | (500 ~ 1 000) V | 1.6 × 10-4
40 ㎐ ~ 1 k㎐
|  | 100 mV | 1.6 × 10-4
|  | 100 mV ~ 100 V | 1.0 × 10-4
|  | (100 ~ 300) V | 1.9 × 10-4
|  | (300 ~ 1 000) V | 1.6 × 10-4
AC Current |  | 40 ㎐
|  | 200 mA | 3.3 × 10-4
|  | 200 mA ~ 2 A | 4.5 × 10-4
|  | (2 ~ 20) A | 6.2 × 10-4
|  | (20 ~ 40) A | 6.5 × 10-4
40 ㎐ ~ 1 k㎐
|  | 200 mA | 3.0 × 10-4
|  | 200 mA ~ 2 A | 4.3 × 10-4
|  | (2 ~ 20) A | 6.0 × 10-4
|  | (20 ~ 40) A | 2.1 × 10-3
Frequency |  | 10 ㎐ | 5.8 × 10-4
|  | 10 ㎐ ~ 1 k㎐ | 1.2 × 10-4
Timer |  | 10 ms | 6.0 × 10-4
|  | (10 ~ 100) ms | 6.0 × 10-5
|  | 100 ms ~ 100 s | 6.0 × 10-6
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF signal generators | 40426 |  |  | Audio Analyzer, Digital
Frequency Test |  | 1 ㎐ ~ 10 M㎐ | 5.8 × 10-9 | Multimeter/
SS-CAL-40426
Output Level Test |  | (40 ㎐)
|  | 1 mV | 4.8 × 10-3
|  | (1 ~ 10) mV | 6.0 × 10-4
|  | 10 mV ~ 100 mV | 1.9 × 10-4
|  | 100 mV ~ 10 V | 1.5 × 10-4
|  | (10 ~ 100) V | 1.4 × 10-4
(40 ㎐ ~ 100 ㎐)
|  | 1 mV | 4.8 × 10-3
|  | (1 ~ 10) mV | 5.8 × 10-3
|  | (10 ~ 100) mV | 1.7 × 10-4
|  | 100 mV ~ 100 V | 1.2 × 10-4
100 ㎐ ~ 1 k㎐
|  | 1 mV | 2.5 × 10-3
|  | (1 ~ 10) mV | 3.6 × 10-4
|  | (10 ~ 100) mV | 1.4 × 10-4
|  | 100 mV ~ 100 V | 9.8 × 10-4
(1 k㎐ ~ 10 k㎐)
|  | 1 mV | 4.8 × 10-3
|  | (1 ~ 10) mV | 5.8 × 10-4
|  | (10 ~ 100) mV | 1.7 × 10-4
|  | 100 mV ~ 100 V | 1.2 × 10-4
(10 k㎐ ~ 100 k㎐)
|  | 1 mV | 2.4 × 10-2
|  | (1 ~ 10) mV | 3.2 × 10-3
|  | (10 ~ 100) mV | 1.0 × 10-3
|  | 100 mV ~ 100 V | 8.2 × 10-4
DC Offset |  | (-20 ~ 20) V | 6.0 × 10-5
Output Level Flatness Test |  | (40 ㎐ ~ 10 k㎐)
|  | 1 k㎐,0 ㏈ Ref | 0.006 ㏈
100 mV ~ 30 V
(100 k㎐)
1 k㎐,0 ㏈ Ref
|  | 100 mV | 0.007 ㏈
|  | 100 mV ~ 1 V | 0.007 ㏈
|  | (1 ~ 30) V | 0.006 ㏈
Attenuator Test |  | (100 ㎐ ~ 100 k㎐)
|  | (10 ~ -60)㏈ | 0.10 ㏈
Distortion |  | 20 ㎐ ~ 100 k㎐
|  | (10.0 ~ 0.010) % | 6.6 × 10-2
Rise/Fall Time |  | 10 μs ~ 1 ns | 1.2 × 10-3
|  | 1 ~ 100 ps | 1.2 × 10-3
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF signal generators | 40426 |  |  | Audio Analyzer, Digital
Duty cycle |  | (1 ~ 99) % | 0.006 2 % | Multimeter/
SS-CAL-40426
Transistor curve tracers | 40432 |  |  | Digital Multimeter/
DC Voltage(Source) |  | (±) |  | SS-CAL-40432
|  | 10 mV | 2.0 × 10-5
|  | (10 ~ 100) mV | 6.0 × 10-6
|  | 100 mV ~ 1 000 V | 2.0 × 10-5
DC Current(Source) |  | (±)
|  | 1 nA | 1.2 × 10-2
|  | (1 ~ 10) nA | 8.2 × 10-3
|  | (10 ~ 100) nA | 7.2 × 10-3
|  | 100 nA ~ 1 μA | 4.8 × 10-4
|  | (1 ~ 10) μA | 6.0 × 10-5
|  | (10 ~ 100) μA | 1.4 × 10-5
|  | 100 μA ~ 10 mA | 2.0 × 10-5
|  | (10 ~ 100) mA | 5.0 × 10-5
|  | 100 mA ~ 1 A | 2.2 × 10-4
|  | (1 ~ 10) A | 5.0 × 10-4
DC Voltage(Measure) |  | (±)
|  | 10 mV | 1.0 × 10-4
|  | (10 ~ 100) mV | 1.0 × 10-5
|  | 100 mV ~ 1 000 V | 6.0 × 10-5
AC/DC high voltage generators | 40434 |  |  | High Voltage Digital Meter/
DC Voltage |  | (±) |  | SS-CAL-40434
|  | 0.1 kV | 6.0 × 10-3
|  | (0.1 ~ 100) kV | 3.0 × 10-3
AC Voltage |  | (50 ~ 60) ㎐
|  | 0.1 kV | 1.2 × 10-2
|  | (0.1 ~ 100) kV | 1.0 × 10-2
AC/DC high voltage probes | 40435 |  |  | DC Power Supply/
DC Voltage |  | (±) |  | SS-CAL-40435
|  | (0 ~ 0.1) kV | 8.0 × 10-3
|  | (0.1 ~ 100) kV | 3.0 × 10-3
AC Voltage |  | (50 ~ 60) ㎐
|  | (0 ~ 0.1) kV | 2.0 × 10-2
|  | (0.1 ~ 100) kV | 1.1 × 10-2
Ratio (DC) |  | (±)
0.1 kV
|  | (100 ~ 100 000 : 1) | 1%
(0.1 ~ 100) kV
|  | (100 ~ 100 000 : 1) | 0.3%
Ratio (AC) |  | (50 ~ 60) ㎐
0.1 kV
|  | (100 ~ 100 000 : 1) | 1.4%
(0.1 ~ 100) kV
|  | (100 ~ 100 000 : 1) | 1.1%
Logic analyzers | 40436 | (±) |  | Calibartor/
DC Voltage |  | 10 mV | 1.0 × 10-4 | SS-CAL-40436
|  | (10 ~ 100) mV | 1.1 × 10-5
|  | 100 mV ~ 10 V | 8.0 × 10-6
501. Contact Thermometry
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Temperature generators; oven, | 50101 |  |  | SPRT, Thermocouple
furnaces, isothermal |  |  |  | /SS-CAL-50101
liquid baths, dry-block, calibrators
|  | 0 ℃ | 0.02 ℃
|  | (-196 ~ -80) ℃ | 0.06 ℃
|  | (-80 ~ 550) ℃ | 0.02 ℃
|  | (550 ~ 1 100) ℃ | 1.3 ℃
|  | (1 100 ~ 1 500) ℃ | 2.5 ℃
Temperature indicators | 50102 |  |  | SPRT, Thermocouple
/recorders/controllers, |  |  |  | /SS-CAL-50102
Temperature calibrators
Thermometric type |  | (-196 ~ 550) ℃ | 0.5 ℃
(with sensor) |  | (550 ~ 1 100) ℃ | 1.3 ℃
|  | (1 100 ~ 1 500) ℃ | 2.5 ℃
Resistance type |  | (-196 ~ 250) ℃ | 0.05 ℃ | SPRT
(with sensor) |  | (250 ~ 550) ℃ | 0.07 ℃ | /SS-CAL-50102
Thermometric type(without sensor) |  | (-196 ~ 1 500) ℃ | 0.6 ℃ | Temperature Calibrator
Resistance type(without sensor) |  | (-196 ~ 550) ℃ | 0.06 ℃ | /SS-CAL-50102-2
Temperature Calibrators |  |  |  | Temperature Calibrator
Resistance type (Source) |  | (-196 ~ 600) ℃ | 0.12 ℃ | /SS-CAL-50102-2
Thermometric type (Source) |  | (-196 ~ 1 500) ℃ | 0.45 ℃
Resistance type (Measure) |  | (-196 ~ 600) ℃ | 0.05 ℃
Thermometric type (Measure) |  | (-196 ~ 1 500) ℃ | 0.44 ℃
Glass thermometers; | 50103 |  |  | SPRT
liquid-in-glass, Beckmann |  |  |  | /SS-CAL-50103
Glass thermometers |  | (-50 ~ 400) ℃ | 0.05 ℃
Resistance thermometers; | 50104 |  |  | SPRT
SPRT, IPRT, thermistors, etc. |  |  |  | /SS-CAL-50104
IPRT |  | (-196 ~ 550) ℃ | 0.04 ℃
Thermal expansion | 50105 |  |  | SPRT
thermometers ; bimetal, |  |  |  | /SS-CAL-50105
gas or liquid typy
bimetal |  | (-50 ~ 0) ℃ | 0.3 ℃
|  | (0 ~ 250) ℃ | 0.6 ℃
|  | (250 ~ 500) ℃ | 1.4 ℃
Thermomecoules | 50106 |  |  | SPRT, Thermocouple
noble metal, base metal, |  |  |  | /SS-CAL-50106
pure metal, special type, etc.
base Thermocouple |  | (-196 ~ 550) ℃ | 0.5 ℃
thermometers |  | (550 ~ 1 100) ℃ | 1.3 ℃
Temperature transducers | 50107 |  |  | SPRT, Thermocouple
|  | (-196 ~ 550) ℃ | 0.46 ℃ | /SS-CAL-50107
|  | (550 ~ 1 100) ℃ | 1.3 ℃
502. Non Contact thermometry
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Radiation thermometers | 50204 |  |  | Standard Radiation
|  | (0 ~ 75) ℃ | 1.0 ℃ | Thermometer
|  | (75 ~ 500) ℃ | 1.2 ℃ | /SS-CAL-50204
|  | (500 ~ 800) ℃ | 1.6 ℃
|  | (800 ~ 1 200) ℃ | 1.8 ℃
502. Non Contact thermometry
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Thermal image apparatus | 50205 |  |  | Standard Radiation
|  | (0 ~ 150) ℃ | 1.2 ℃ | Thermometer
|  | (150 ~ 300) ℃ | 1.4 ℃ | /SS-CAL-50205
|  | (300 ~ 800) ℃ | 1.6 ℃
|  | (800 ~ 1 200) ℃ | 2.0 ℃
Blackbody furnaces | 50206 |  |  | Standard Radiation
|  | (0 ~ 75) ℃ | 1.1 ℃ | Thermometer
|  | (75 ~ 500) ℃ | 1.3 ℃ | /SS-CAL-50206
|  | (500 ~ 1 200) ℃ | 1.9 ℃
Others Thermometer; | 50207 |  |  | SPRT
ear thermometers, etc. |  |  |  | /SS-CAL-50207
|  | (34 ~ 42) ℃ | 0.1 ℃
503. Humidity
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Relative humidity hygrometers; | 50302 |  |  | Dewpoint Meter
polimer thinfilm, hair, etc. |  | (10 ~ 50) % R.H. | 1.6 % R.H. | /SS-CAL-50302
|  | (50 ~ 70) % R.H. | 2.2 % R.H.
|  | (70 ~ 95) % R.H. | 2.7 % R.H.
|  | (-40 ~ 80) ℃ | 0.6 ℃
Psychrometers; assmann | 50303 |  |  | Dewpoint Meter
ventilated, PRT type, etc. |  |  |  | /SS-CAL-50303
|  | (10 ~ 50) % R.H. | 1.6 % R.H.
|  | (50 ~ 70) % R.H. | 2.2 % R.H.
|  | (70 ~ 95) % R.H. | 2.7 % R.H.
|  | (-40 ~ 60) ℃ | 0.6 ℃
Temperature humidity recorders; | 50304 |  |  | Dewpoint Meter
|  | (10 ~ 50) % R.H. | 1.8 % R.H. | /SS-CAL-50304
|  | (50 ~ 70) % R.H. | 2.3 % R.H.
|  | (70 ~ 95) % R.H. | 2.8 % R.H.
|  | (-20 ~ 50) ℃ | 0.7 ℃
Dew-point/Relative humidity | 50305 |  |  | Dewpoint Meter
Transducers |  |  |  | /SS-CAL-50305
Relative humidity Transducers |  | (10 ~ 50) % R.H. | 1.7 % R.H.
|  | (50 ~ 70) % R.H. | 2.3 % R.H.
|  | (70 ~ 95) % R.H. | 2.8 % R.H.
|  | (-20 ~ 60) ℃ | 0.7 ℃
Humidity generators; two-pressure | 50306 |  |  | Dewpoint Meter
two-temperature, flow mixing |  |  |  | /SS-CAL-50306
humidity gererator, constant temperature
and humidity chamber, etc.
humidity chamber |  | (10 ~ 50) % R.H. | 2.0 % R.H.
|  | (50 ~ 95) % R.H. | 3.0 % R.H.
|  | (-75 ~ 180) ℃ | 0.9 ℃
601. Sound in air
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Sound level meters | 60106 |  |  | Multifunction Acoustic
|  | 125 ㎐ | 0.3 ㏈ | Calibrator/
|  | 250 ㎐ | 0.2 ㏈ | SS-CAL-60106
|  | 500 ㎐ | 0.2 ㏈
|  | 1 k㎐ | 0.2 ㏈
|  | 2 k㎐ | 0.2 ㏈
|  | 4 k㎐ | 0.3 ㏈
|  | 8 k㎐ | 0.3 ㏈
603. Vibration
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Vibration calibrators | 60301 |  |  | Standard Accelerometer/
|  | (20 ~ 160) ㎐ | 1.6 × 10-2 | SS-CAL-60301
|  | (160 ~ 1 250) ㎐ | 1.8 × 10-2
Vibration transducers | 60302 |  |  | Standard Accelerometer/
|  | (10 ~ 160) ㎐ | 1.5 × 10-2 | SS-CAL-60302
|  | (160 ~ 5 000) ㎐ | 1.7 × 10-2
Vibration measuring instruments | 60303 |  |  | Standard Accelerometer/
Acceleration |  | (10 ~ 160) ㎐ | 1.5 × 10-2 | SS-CAL-60303
|  | (160 ~ 2 500) ㎐ | 1.7 × 10-2
Velocity |  | (10 ~ 160) ㎐ | 1.5 × 10-2
|  | (160 ~ 2 500) ㎐ | 1.7 × 10-2
Displacement |  | (10 ~ 160) ㎐ | 1.4 × 10-2
|  | (160 ~ 315) ㎐ | 2.3 × 10-2
|  | (315 ~ 630) ㎐ | 6.6 × 10-2
701. Photometry
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Illuminance meters | 70101 |  |  | Illuminance meter
Illuminance |  | 0.5 lx | 2.7% | Standard lamp
|  | (0.5 ~ 1) lx | 2.3% | SS-CAL-70101
|  | (1 ~ 5) lx | 2.2%
|  | (5 ~ 10) lx | 2.4%
|  | (10 ~ 500) lx | 2.1%
|  | (500 ~ 1 000) lx | 2.2%
|  | (1 000 ~ 3 000) lx | 2.3%
|  | (3 000 ~ 5 000) lx | 2.4%
|  | (5 000 ~ 10 000) lx | 2.7%
|  | (10 000 ~ 20 000) lx | 3.2%
704. Fiber optics
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Broadband light sources | 70402 |  |  | Optical spectrum analyzer
Wavelength |  | (633 ~ 1 640) nm | 0.059 | Optical power meter
SS-CAL-70402
Output ~Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
Optical attenuators | 70410 |  |  | Optical power stabilized and LD
Attenuation |  | 1 310 nm, 1 550 nm |  | Optical power meter
|  | (0 ~ 55) dB | 0.08 | SS-CAL-70410
Fiber-optic power meters | 70412 |  |  | Optical power stabilized and LD
Absolute Optical Power |  | 1 310 nm, 1 550 nm |  | Optical power meter
|  | (-55 ~ 0) dBm | 0.08 | Optical attenuator
SS-CAL-70412
Linearity |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dB | 0.08
Optical loss testers | 70413 |  |  | Optical power stabilized and LD
Wavelength |  | (633 ~ 1 640) nm | 0.059 | Optical power meter
Optical attenuator
Output Optical Power |  | 1 310 nm, 1 550 nm |  | Optical spectrum analyzer
|  | (-55 ~ 0) dBm | 0.08 | SS-CAL-70413
Input Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
Linearity |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dB | 0.08
Optical multimeters | 70415 |  |  | Optical power stabilized and LD
Wavelength |  | (633 ~ 1 640) nm | 0.059 | Optical power meter
Optical attenuator
Output Optical Power |  | 1 310 nm, 1 550 nm |  | Optical spectrum analyzer
|  | (-55 ~ 0) dBm | 0.08 | SS-CAL-70415
Input Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
Linearity |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dB | 0.08
Optical spectrum analyzers | 70417 |  |  | Optical power stabilized and LD
Wavelength |  | 1 310 nm, 1 550 nm | 0.017 | Optical power meter
Optical attenuator
Resolution |  | 1 310 nm, 1 550 nm |  | SS-CAL-70417
|  | (0.1 ~ 1) nm | 0.017
Input Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
Linearity |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dB | 0.08
704. Fiber optics
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Optical time domain reflectometers : | 70418 |  |  | Optical spectrum analyzer
OTDR |  |  |  | Standard Fiber
Wavelength |  | 1 310 nm, 1 550 nm | 0.059 | SS-CAL-70418
Length |  | 1 310 nm
|  | 3 km | 0.08 m
|  | 13 km | 0.32 m
1 550 nm
|  | 3 km | 0.08 m
|  | 13 km | 0.32 m
Loss |  | 1 310 nm
|  | 1.01 dB | 0.04 dB
|  | 4.24 dB | 0.11 dB
1 550 nm
|  | 0.56 dB | 0.04 dB
|  | 2.44 dB | 0.07 dB
Multi-laser wavelength meters | 70426 |  |  | Frequency stabilized laser and LD
Wavelength |  | 1 310 nm | 5 | Optical power meter
|  | (1 460 ~ 1 580) nm | 2 | Optical attenuator
Multi-Wavelengh Meter
Input Optical Power |  | 1 310 nm, 1 550 nm |  | SS-CAL-70426
|  | (-55 ~ 0) dBm | 0.08
Frequency stabilized lasers and LDs | 70429 |  |  | Multi-Wavelength Meter
Wavelength |  | 1 310 nm | 5 | Optical power meter
|  | (1 460 ~ 1 640) nm | 5 | SS-CAL-70429
Output Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
ASE light sources | 70430 |  |  | Optical spectrum analyzer
Wavelength |  | (633 ~ 1 640) nm | 0.059 | Optical power meter
SS-CAL-70430
Output Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
CW-laser wavelength meters | 70431 |  |  | Frequency stabilized laser and LD
Wavelength |  | 1 310 nm | 5 | Optical power meter
|  | (1 460 ~ 1 580) nm | 2 | Optical attenuator
Multi-Wavelengh Meter
Input Optical Power |  | 1 310 nm, 1 550 nm |  | SS-CAL-70431
|  | (-55 ~ 0) dBm | 0.08
Optical power stabilized and LDs | 70433 |  |  | Optical spectrum analyzer
Wavelength |  | (633 ~ 1 640) nm | 0.059 | Optical power meter
SS-CAL-70433
Output Optical Power |  | 1 310 nm, 1 550 nm
|  | (-55 ~ 0) dBm | 0.08
901. Chemical analysis
Measured Quantity Instrument or Gauge | Field Code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Gas analyzers | 90103 |  |  | GAS CRM
Oxygen(O2) |  | (0 ~ 18.0) cmol/mol | 2.2 × 10-2 | /SS-CAL-90103
|  | (18.1 ~ 21.0) cmol/mol | 1.9 × 10-2
Carbon monoxide(CO) |  | (0 ~ 75.0) µmol/mol | 2.0 × 10-2
|  | (75.1 ~ 150.0) µmol/mol | 1.9 × 10-2
Hydrogen sulfide(H2S) |  | (0 ~ 25.0) µmol/mol | 3.9 × 10-2
|  | (25.1 ~ 100.0) µmol/mol | 3.5 × 10-2
Methane(CH4) |  | (0 ~ 1.25) cmol/mol | 2.4 × 10-2
|  | (1.26 ~ 2.50) cmol/mol | 2.0 × 10-2
Carbon dioxide(CO2) |  | (0 ~ 2 500) µmol/mol | 2.0 × 10-2
|  | (0.26 ~ 2.50) cmol/mol | 2.0 × 10-2
Hydrogen(H2) |  | (0 ~ 200) µmol/mol | 1.0 × 10-2
|  | (201 ~ 1 000) µmol/mol | 1.6 × 10-2
|  | (0.10 ~ 2.00) cmol/mol | 1.5 × 10-2
√("(0.7 μm)2" +(3×10－6×𝑙0)2)
√("(0.37 μm)2" +(45×10－6×𝑙0)2)
√("(6.2 μm)2" +(7×10－6×𝑙0)2)
√("(1.6 μm)2" +(3.4×10－6×𝑙0)) 2
√("(7.3 μm)2" +(7×10－6×𝑙0)2)
√("(0.9 μm)2" +(1.9×10－6×𝑙0)2)
√("(0.8 μm)2" +(7×10－6×𝑙0)2)
√("(0.85 μm)2" +(14×10－6×𝑙0)2)
√("(1.1 μm)2" +(7.2×10－6×𝑙0)2)
√("(1.1 μm)2" +(6.4×10－6×𝑙0)2)
√("(0.66 μm)2" +(6.0×10－6×𝑙0)2)
√("(1.1 μm)2" +(5.2×10－6×𝑙0)2)
√("(0.71 μm)2" +(14×10－6×𝑙0)2)
√("(0.32 μm)2" +(45×10－6×𝑙0)2)
√("(1.5 μm)2" +(1.4×10－6×𝑙0)2)
√("(1.5 μm)2" +(2×10－6×𝑙0)2)
√("(0.82 μm)2" +(6.8×10－6×𝑙0)2)